## log
FPA SN 102 test log
Date and Time: 4/14/2022 12:30 PM
Test IBAT 677610
Test plan JPL D-107025
Parts
FPA 10501725-1 Rev B SN102, assembly IBAT 675812
with Detector PREFIRE6.4, K210104.2 (2,-1)
with PWA SN0004
Unwired pixels (with pixel numbering 1-64)
Dark pixels: for all rows, pixels 3, 4, 9, 10, 18, 19, 36, 37
Bad pixels from resistance measurement: A 39, 64; C 33-35, 38-64; E 40, 44, 47, 48, 53, 55, 58, 60, 62-64; G 38, 41, 48, 60; H 41, 51, 53, 58, 60, 61
Corresponding ROIC channels shorted
Setup
Chamber KJLC A0058497
CDH (PWB Rev A, PWA Rev A, SN 001) with GSE library v1.2.0
Modified BFE power board powered by E3631A
Copper baffle, copper bracket, and copper support
Setttings
Power supply (E3631A) 6.001 V 632 mA, 3.30 V 136 mA
Chamber pressure < 0.26 mTorr
Chopper frequency, sclk_divider, window, and integration: 62.5 kHz, 0x01, 0, 194 ms
Nominal detectivity is calculated with 1 Hz ≤ noise frequency < FPS/2.
Blackbody temperature: 1000 ± .1 K + source calibration accuracy
Shutter drive frequency: 50 mHz
Aperture diameters: 0.2", 0.1" and 0.05"
Measured average FPS = 4.983 Hz
Data file folder names
Noise data: 20220414.144109, 20220414.145935, 20220414.162508, 20220414.164046
Response data: 20220414.152204 for 0.2", 20220414.152744 for 0.1", 20220414.153315 for 0.05"

## power
Optical power calculation
Blackbody temperature (K) | 1000
Exitance (W/m^2) | 56700
ZnSe window transmission | 0.7
Aperture-detector distance (in) | 7.74
Detector area (m^2) | 3.24e-08
Aperture diameter (in) | 0.2 | 0.1 | 0.05
Incidance (W/m^2) | 6.6241 | 1.6562 | 0.41407
Optical power (W) | 2.1462e-07 | 5.3662e-08 | 1.3416e-08

## pixels
Pixel number | Response (counts) for aperture (in) |  |  | Gain (10^9 counts/W) for aperature (in) |  |  |  | Noise spectral density (counts/rtHz) between 1 Hz and FPS/2, 4 runs |  |  |  |  | Detectivity (10^8 Jones)
| 0.2 | 0.1 | 0.05 | 0.2 | 0.1 | 0.05 | average | 1st | 2nd | 3rd | 4th | rms avg
A1 | 3270.3 | 784.856 | 197.113 | 15.2376 | 14.6259 | 14.6925 | 14.852 | 0.72308 | 0.69777 | 0.70463 | 0.70515 | 0.70772 | 3.7774 | Legend
A2 | 3255.12 | 808.353 | 203.193 | 15.1668 | 15.0638 | 15.1457 | 15.1255 | 0.73766 | 0.74215 | 0.72455 | 0.79485 | 0.75028 | 3.6288 | Pixel 1-2
A3 | 31.008 | 4.992 | 0.996 | 0.14448 | 0.093027 | 0.07424 | 0.10392 | 0.58657 | 0.5926 | 0.54833 | 0.62224 | 0.58803 | 0.031809 | Pixel 5-64 with D* > 2e8 Jones
A4 | 30.996 | 4.004 | 0 | 0.14442 | 0.074615 | 0 | 0.073013 | 0.64296 | 0.63799 | 0.59454 | 0.67003 | 0.63695 | 0.020633 | Connected pixel 5-64 with D* < 0.8e8 Jones
A5 | 6759.611 | 1742.618 | 431.834 | 31.4956 | 32.474 | 32.1882 | 32.0526 | 0.74377 | 0.7397 | 0.76991 | 0.73735 | 0.7478 | 7.7153 | Dark pixel
A6 | 6694.541 | 1725.805 | 426.735 | 31.1924 | 32.1607 | 31.8082 | 31.7204 | 0.72806 | 0.74982 | 0.69741 | 0.69895 | 0.71889 | 7.9424 | Bad pixel from resistance measurement
A7 | 6728.736 | 1753.29 | 431.834 | 31.3518 | 32.6729 | 32.1882 | 32.071 | 0.76776 | 0.7518 | 0.781 | 0.76856 | 0.76735 | 7.523
A8 | 6902.315 | 1794.87 | 441.98 | 32.1605 | 33.4478 | 32.9445 | 32.8509 | 0.81654 | 0.87607 | 0.82739 | 0.84503 | 0.84156 | 7.0265
A9 | 30.969 | 5.004 | 1 | 0.1443 | 0.093251 | 0.074538 | 0.10403 | 0.65619 | 0.67267 | 0.6198 | 0.69444 | 0.66134 | 0.028314
A10 | 33.978 | 2.002 | 0 | 0.15832 | 0.037308 | 0 | 0.065208 | 0.64188 | 0.64689 | 0.60123 | 0.69002 | 0.64577 | 0.018176
A11 | 6903.22 | 1816.932 | 447.336 | 32.1647 | 33.8589 | 33.3437 | 33.1224 | 0.69237 | 0.69876 | 0.73094 | 0.69969 | 0.7056 | 8.4496
A12 | 6821.984 | 1833.606 | 450.145 | 31.7862 | 34.1696 | 33.5531 | 33.1696 | 0.72036 | 0.75746 | 0.73472 | 0.77736 | 0.74779 | 7.9842
A13 | 6956.82 | 1832.49 | 448.902 | 32.4145 | 34.1488 | 33.4605 | 33.3412 | 0.74791 | 0.77133 | 0.76967 | 0.72271 | 0.75316 | 7.9683
A14 | 6988.274 | 1846.572 | 454.256 | 32.561 | 34.4112 | 33.8595 | 33.6106 | 0.73018 | 0.71176 | 0.71318 | 0.74127 | 0.7242 | 8.3539
A15 | 6853.756 | 1811.848 | 447.265 | 31.9343 | 33.7641 | 33.3384 | 33.0123 | 0.68842 | 0.67706 | 0.66015 | 0.70852 | 0.68377 | 8.6904
A16 | 6075.212 | 1856.25 | 456.259 | 28.3067 | 34.5916 | 34.0088 | 32.3024 | 0.73273 | 0.73954 | 0.7523 | 0.7652 | 0.74755 | 7.778
A17 | 7015.14 | 1842.935 | 453.924 | 32.6862 | 34.3435 | 33.8348 | 33.6215 | 0.7841 | 0.83219 | 0.81878 | 0.8148 | 0.81266 | 7.447
A18 | 54.044 | 0 | 1.998 | 0.25181 | 0 | 0.14893 | 0.13358 | 0.6649 | 0.64237 | 0.59079 | 0.69639 | 0.64975 | 0.037005
A19 | 41.031 | 1 | 1.001 | 0.19118 | 0.018635 | 0.074613 | 0.094809 | 0.59604 | 0.60962 | 0.61803 | 0.68039 | 0.62685 | 0.027224
A20 | 6289.47 | 1819.76 | 450.214 | 29.305 | 33.9116 | 33.5583 | 32.2583 | 0.73714 | 0.76617 | 0.74756 | 0.77208 | 0.75586 | 7.6819
A21 | 6921.022 | 1825.95 | 451.26 | 32.2477 | 34.0269 | 33.6362 | 33.3036 | 0.73444 | 0.73403 | 0.75844 | 0.74799 | 0.74379 | 8.0596
A22 | 5717.556 | 1838.551 | 452.91 | 26.6403 | 34.2618 | 33.7592 | 31.5537 | 0.80108 | 0.77485 | 0.77772 | 0.80407 | 0.78954 | 7.1936
A23 | 7022.368 | 1843.608 | 455.282 | 32.7199 | 34.356 | 33.936 | 33.6706 | 0.737 | 0.77436 | 0.71137 | 0.75372 | 0.74447 | 8.141
A24 | 6835.872 | 1843.608 | 454.203 | 31.8509 | 34.356 | 33.8556 | 33.3542 | 0.70474 | 0.72622 | 0.73406 | 0.70699 | 0.71811 | 8.3605
A25 | 7095.086 | 1870.365 | 461.916 | 33.0587 | 34.8546 | 34.4305 | 34.1146 | 0.79671 | 0.82186 | 0.80324 | 0.82703 | 0.81231 | 7.5595
A26 | 6377.085 | 1824.57 | 446.355 | 29.7133 | 34.0012 | 33.2706 | 32.3284 | 0.69953 | 0.72184 | 0.74458 | 0.68173 | 0.71231 | 8.1693
A27 | 7131.96 | 1870.199 | 458.127 | 33.2305 | 34.8515 | 34.1481 | 34.0767 | 0.68036 | 0.71395 | 0.68642 | 0.67603 | 0.68935 | 8.898
A28 | 5979.376 | 1889.91 | 463.125 | 27.8602 | 35.2188 | 34.5206 | 32.5332 | 1.0666 | 1.1223 | 1.0833 | 1.1163 | 1.0974 | 5.3364
A29 | 7298.82 | 1938.456 | 476.773 | 34.008 | 36.1235 | 35.5379 | 35.2231 | 0.84698 | 0.87595 | 0.84225 | 0.85539 | 0.85524 | 7.4133
A30 | 6704.28 | 1942.408 | 476.755 | 31.2378 | 36.1972 | 35.5366 | 34.3238 | 0.83556 | 0.84743 | 0.82556 | 0.88949 | 0.84986 | 7.2698
A31 | 7254.72 | 1939.462 | 470.92 | 33.8025 | 36.1423 | 35.1016 | 35.0155 | 0.77506 | 0.73605 | 0.76277 | 0.76821 | 0.76067 | 8.2858
A32 | 6412.12 | 1886.092 | 456.806 | 29.8765 | 35.1477 | 34.0496 | 33.0246 | 0.78114 | 0.77971 | 0.77124 | 0.81726 | 0.78754 | 7.5481
A33 | 6917.13 | 1864.356 | 450.214 | 32.2295 | 34.7426 | 33.5583 | 33.5101 | 0.8132 | 0.84392 | 0.87626 | 0.86434 | 0.84976 | 7.0982
A34 | 6649.83 | 1862.38 | 447.066 | 30.9841 | 34.7058 | 33.3236 | 33.0045 | 0.73505 | 0.74207 | 0.77749 | 0.76831 | 0.75594 | 7.8589
A35 | 6681.51 | 1824.705 | 438.418 | 31.1317 | 34.0037 | 32.679 | 32.6048 | 0.78602 | 0.75852 | 0.74251 | 0.79477 | 0.77074 | 7.6146
A36 | 57.013 | 1.001 | 1.008 | 0.26565 | 0.018654 | 0.075135 | 0.11981 | 0.73238 | 0.70443 | 0.71522 | 0.8146 | 0.74292 | 0.029029
A37 | 52.018 | 0 | 0.999 | 0.24237 | 0 | 0.074464 | 0.10561 | 0.67428 | 0.68319 | 0.68113 | 0.7885 | 0.70836 | 0.026837
A38 | 6667.448 | 1891.819 | 447.265 | 31.0662 | 35.2544 | 33.3384 | 33.2197 | 0.75965 | 0.7908 | 0.76932 | 0.77912 | 0.77481 | 7.7174
A39 | 45.03 | 3.003 | 0 | 0.20981 | 0.055961 | 0 | 0.088591 | 0.74545 | 0.73328 | 0.76472 | 0.82845 | 0.76885 | 0.020741
A40 | 141.102 | 0.99 | 3.003 | 0.65745 | 0.018449 | 0.22384 | 0.29991 | 1.3283 | 1.3132 | 1.3102 | 1.9246 | 1.4925 | 0.036171
A41 | 6724.08 | 1873.981 | 444.92 | 31.3301 | 34.922 | 33.1636 | 33.1386 | 0.8164 | 0.83678 | 0.82275 | 0.85923 | 0.83395 | 7.1526
A42 | 6396.39 | 1800.81 | 423.89 | 29.8032 | 33.5584 | 31.5961 | 31.6526 | 0.82159 | 0.79608 | 0.81494 | 0.81612 | 0.81224 | 7.0145
A43 | 6467.266 | 1784.328 | 422.064 | 30.1335 | 33.2513 | 31.46 | 31.6149 | 0.84803 | 0.86674 | 0.84306 | 0.87832 | 0.85916 | 6.6236
A44 | 6336.044 | 1849.536 | 438.892 | 29.522 | 34.4665 | 32.7143 | 32.2343 | 0.85964 | 0.86926 | 0.88341 | 0.8956 | 0.87708 | 6.6153
A45 | 6740.91 | 1861.2 | 442.65 | 31.4085 | 34.6838 | 32.9944 | 33.0289 | 0.90007 | 0.88926 | 0.93155 | 0.93931 | 0.91529 | 6.4955
A46 | 6631.02 | 1836.573 | 440.834 | 30.8965 | 34.2249 | 32.8591 | 32.6601 | 0.92604 | 0.90642 | 0.97953 | 0.97701 | 0.94778 | 6.2027
A47 | 6513.21 | 1790.09 | 428.824 | 30.3475 | 33.3587 | 31.9639 | 31.89 | 0.973 | 0.91124 | 0.94983 | 0.96095 | 0.94904 | 6.0485
A48 | 5714.106 | 1558.055 | 375.192 | 26.6242 | 29.0347 | 27.9662 | 27.875 | 0.89628 | 0.87372 | 0.8736 | 0.91565 | 0.88999 | 5.6377
A49 | 502.85 | 136.032 | 34.986 | 2.343 | 2.535 | 2.6078 | 2.4953 | 0.97352 | 0.97608 | 0.98082 | 1.0182 | 0.98733 | 0.45491
A50 | 980.712 | 105.966 | 22.983 | 4.5695 | 1.9747 | 1.7131 | 2.7524 | 2.661 | 2.5594 | 2.1054 | 4.3185 | 3.0296 | 0.16353
A51 | 5640.03 | 1522.941 | 369.764 | 26.2791 | 28.3803 | 27.5616 | 27.407 | 0.93489 | 0.9019 | 0.95128 | 1.0114 | 0.95071 | 5.189
A52 | 6444.473 | 1744.596 | 420.831 | 30.0273 | 32.5109 | 31.3681 | 31.3021 | 0.90284 | 0.90907 | 0.96891 | 0.97939 | 0.94068 | 5.9897
A53 | 6359.247 | 1701.08 | 418.07 | 29.6302 | 31.7 | 31.1623 | 30.8308 | 1.0067 | 0.9589 | 1.0048 | 1.0568 | 1.0074 | 5.5089
A54 | 5532.12 | 1712.682 | 421.95 | 25.7763 | 31.9162 | 31.4515 | 29.7146 | 1.0962 | 1.0781 | 1.1137 | 1.1098 | 1.0995 | 4.8645
A55 | 6619.616 | 1728.237 | 431.925 | 30.8433 | 32.206 | 32.195 | 31.7481 | 1.026 | 1.0189 | 1.0187 | 1.0779 | 1.0357 | 5.5178
A56 | 5389.02 | 1737.45 | 435.15 | 25.1095 | 32.3777 | 32.4354 | 29.9742 | 0.98455 | 0.95086 | 0.98048 | 0.9816 | 0.97447 | 5.5367
A57 | 6403.842 | 1683.822 | 424.242 | 29.8379 | 31.3784 | 31.6223 | 30.9462 | 1.0112 | 1.0275 | 1.0701 | 1.0458 | 1.0389 | 5.3619
A58 | 6391.907 | 1732.728 | 438.165 | 29.7823 | 32.2897 | 32.6601 | 31.5774 | 1.0266 | 1.0185 | 1.0768 | 1.0505 | 1.0433 | 5.4478
A59 | 6595.808 | 1736.133 | 438.165 | 30.7324 | 32.3532 | 32.6601 | 31.9152 | 1.0229 | 0.99504 | 1.0149 | 1.0395 | 1.0182 | 5.6419
A60 | 6783.48 | 1794.046 | 452.952 | 31.6068 | 33.4324 | 33.7623 | 32.9339 | 0.95862 | 0.94167 | 0.98481 | 1.0259 | 0.97827 | 6.0598
A61 | 7066.821 | 1858.521 | 473.841 | 32.927 | 34.6339 | 35.3194 | 34.2934 | 0.99278 | 1.0241 | 1.015 | 1.0506 | 1.0208 | 6.0469
A62 | 7220.07 | 1988.715 | 505.891 | 33.6411 | 37.0601 | 37.7083 | 36.1365 | 1.067 | 1.0393 | 1.0684 | 1.0925 | 1.0669 | 6.0964
A63 | 7908.224 | 2102.614 | 537.38 | 36.8474 | 39.1826 | 40.0555 | 38.6952 | 1.095 | 1.1117 | 1.0818 | 1.0971 | 1.0964 | 6.3525
A64 | 10.005 | 13 | 2 | 0.046617 | 0.24226 | 0.14908 | 0.14598 | 0.92414 | 0.86016 | 0.91159 | 1.0451 | 0.93771 | 0.028023
B1 | 2366.508 | 575.055 | 143.032 | 11.0265 | 10.7163 | 10.6614 | 10.8014 | 0.64349 | 0.67539 | 0.61537 | 0.66069 | 0.64912 | 2.9952
B2 | 2734.585 | 669.494 | 168.08 | 12.7415 | 12.4762 | 12.5284 | 12.582 | 0.64471 | 0.63045 | 0.63507 | 0.63482 | 0.63628 | 3.5594
B3 | 1.001 | 12 | 2.004 | 0.004664 | 0.22362 | 0.14938 | 0.12589 | 0.63244 | 0.60084 | 0.53594 | 0.60683 | 0.59508 | 0.038079
B4 | 3.003 | 13 | 0.999 | 0.013992 | 0.24226 | 0.074464 | 0.11024 | 0.75615 | 0.70722 | 0.62974 | 0.70196 | 0.70022 | 0.028338
B5 | 5809.44 | 1512.084 | 381.024 | 27.0684 | 28.178 | 28.4009 | 27.8824 | 0.87115 | 0.84189 | 0.7451 | 0.83865 | 0.82556 | 6.0793
B6 | 5995.866 | 1516.9 | 376.941 | 27.937 | 28.2677 | 28.0966 | 28.1005 | 0.76903 | 0.79134 | 0.69596 | 0.84386 | 0.77687 | 6.5109
B7 | 5701.18 | 1495.143 | 373.808 | 26.564 | 27.8623 | 27.8631 | 27.4298 | 0.94879 | 0.97783 | 0.84578 | 0.95615 | 0.93353 | 5.2889
B8 | 6019.713 | 1532.811 | 377.33 | 28.0481 | 28.5642 | 28.1256 | 28.246 | 0.88859 | 0.86639 | 0.80883 | 0.86537 | 0.8578 | 5.9271
B9 | 6.006 | 10.998 | 2 | 0.027984 | 0.20495 | 0.14908 | 0.12734 | 0.83847 | 0.85978 | 0.75766 | 0.8984 | 0.84015 | 0.027282
B10 | 10.998 | 10.998 | 1.988 | 0.051244 | 0.20495 | 0.14818 | 0.13479 | 0.71218 | 0.71486 | 0.6815 | 0.74633 | 0.71409 | 0.033977
B11 | 5779.512 | 1525.902 | 374.016 | 26.929 | 28.4355 | 27.8786 | 27.7477 | 0.82537 | 0.80744 | 0.80474 | 0.83453 | 0.81811 | 6.105
B12 | 6011.642 | 1549.8 | 381.808 | 28.0105 | 28.8808 | 28.4594 | 28.4502 | 0.81301 | 0.85656 | 0.77894 | 0.8476 | 0.8246 | 6.2103
B13 | 5705.332 | 1515.226 | 371.041 | 26.5833 | 28.2365 | 27.6568 | 27.4922 | 0.8099 | 0.77272 | 0.71872 | 0.77378 | 0.76947 | 6.4312
B14 | 5976.452 | 1541.928 | 377.136 | 27.8466 | 28.7341 | 28.1111 | 28.2306 | 0.79898 | 0.8004 | 0.76719 | 0.79698 | 0.791 | 6.4241
B15 | 5799.692 | 1539.776 | 378.886 | 27.023 | 28.694 | 28.2416 | 27.9862 | 0.72383 | 0.73607 | 0.68858 | 0.72767 | 0.71926 | 7.0037
B16 | 6010.822 | 1561.608 | 382.932 | 28.0067 | 29.1009 | 28.5432 | 28.5502 | 0.82149 | 0.77529 | 0.76759 | 0.78384 | 0.78733 | 6.5272
B17 | 5736.788 | 1527.82 | 375.964 | 26.7299 | 28.4712 | 28.0238 | 27.7416 | 0.85148 | 0.85388 | 0.79861 | 0.87358 | 0.84484 | 5.9106
B18 | 28.016 | 5.005 | 1.001 | 0.13054 | 0.093269 | 0.074613 | 0.099473 | 0.69015 | 0.66695 | 0.63452 | 0.71748 | 0.67796 | 0.02641
B19 | 21.978 | 7.008 | 0 | 0.1024 | 0.1306 | 0 | 0.077666 | 0.73443 | 0.75062 | 0.66826 | 0.72796 | 0.72099 | 0.01939
B20 | 5841.918 | 1515.08 | 369.978 | 27.2197 | 28.2338 | 27.5776 | 27.677 | 0.87177 | 0.84141 | 0.82853 | 0.84608 | 0.8471 | 5.8811
B21 | 5578.342 | 1490.038 | 367.952 | 25.9916 | 27.7671 | 27.4266 | 27.0618 | 0.84801 | 0.78804 | 0.82825 | 0.83008 | 0.82389 | 5.9124
B22 | 5821.296 | 1517.19 | 372.999 | 27.1236 | 28.2731 | 27.8028 | 27.7332 | 0.7494 | 0.73647 | 0.72924 | 0.74931 | 0.74115 | 6.7354
B23 | 5533.821 | 1478.29 | 365.096 | 25.7842 | 27.5482 | 27.2137 | 26.8487 | 0.7475 | 0.75816 | 0.76047 | 0.75549 | 0.75542 | 6.3975
B24 | 5824.197 | 1512.966 | 367.872 | 27.1372 | 28.1944 | 27.4206 | 27.5841 | 0.78258 | 0.78674 | 0.74656 | 0.79056 | 0.77681 | 6.3917
B25 | 5517.644 | 1478.201 | 363.072 | 25.7088 | 27.5466 | 27.0628 | 26.7727 | 0.83773 | 0.8021 | 0.76757 | 0.81915 | 0.80705 | 5.9712
B26 | 5762.376 | 1498.16 | 363.84 | 26.8491 | 27.9185 | 27.1201 | 27.2959 | 0.76485 | 0.69949 | 0.71498 | 0.75155 | 0.73319 | 6.7012
B27 | 5568.732 | 1492.856 | 365.313 | 25.9468 | 27.8197 | 27.2299 | 26.9988 | 0.85879 | 0.82661 | 0.80706 | 0.85749 | 0.83777 | 5.8009
B28 | 6032.304 | 1582.74 | 384.866 | 28.1068 | 29.4947 | 28.6873 | 28.7629 | 0.73486 | 0.74516 | 0.75842 | 0.76076 | 0.74987 | 6.9043
B29 | 5728.151 | 1565.85 | 383.04 | 26.6896 | 29.1799 | 28.5512 | 28.1403 | 0.80876 | 0.8215 | 0.78956 | 0.82423 | 0.81113 | 6.2447
B30 | 5939.183 | 1582.74 | 387.03 | 27.6729 | 29.4947 | 28.8486 | 28.6721 | 0.80186 | 0.80537 | 0.78689 | 0.7721 | 0.79167 | 6.5191
B31 | 5517.119 | 1498.65 | 366.067 | 25.7064 | 27.9276 | 27.2861 | 26.9734 | 0.74976 | 0.71982 | 0.69545 | 0.70799 | 0.71854 | 6.757
B32 | 5809.242 | 1535.844 | 371.304 | 27.0675 | 28.6208 | 27.6764 | 27.7882 | 0.80136 | 0.79641 | 0.75821 | 0.78028 | 0.78425 | 6.3779
B33 | 5577.84 | 1513.015 | 369.024 | 25.9893 | 28.1953 | 27.5065 | 27.2304 | 1.0078 | 0.98262 | 1.002 | 1.0594 | 1.0133 | 4.837
B34 | 5848.192 | 1584.048 | 383.105 | 27.249 | 29.519 | 28.5561 | 28.4414 | 1.0464 | 1.0077 | 1.0749 | 1.1077 | 1.0598 | 4.8305
B35 | 5491.952 | 1498.012 | 364.977 | 25.5891 | 27.9157 | 27.2048 | 26.9032 | 0.95786 | 0.94675 | 0.94427 | 1.0347 | 0.97161 | 4.9841
B36 | 5.985 | 9 | 3.009 | 0.027886 | 0.16772 | 0.22429 | 0.13996 | 1.0177 | 0.9685 | 1.009 | 1.118 | 1.0298 | 0.024465
B37 | 6.006 | 11.016 | 0 | 0.027984 | 0.20529 | 0 | 0.077757 | 1.1661 | 1.1301 | 1.1551 | 1.2705 | 1.1817 | 0.011844
B38 | 5858.775 | 1600.64 | 384.312 | 27.2983 | 29.8282 | 28.646 | 28.5908 | 0.97876 | 1.0105 | 0.96182 | 1.0055 | 0.98934 | 5.2018
B39 | 5724.108 | 1577.22 | 383.04 | 26.6708 | 29.3918 | 28.5512 | 28.2046 | 1.1824 | 1.1491 | 1.1769 | 1.2131 | 1.1806 | 4.3002
B40 | 5767.125 | 1584.048 | 378.805 | 26.8712 | 29.519 | 28.2355 | 28.2086 | 1.0368 | 0.98573 | 1.0488 | 1.0134 | 1.0215 | 4.9708
B41 | 5608.2 | 1538.064 | 368.988 | 26.1307 | 28.6621 | 27.5038 | 27.4322 | 1.0709 | 1.0445 | 1.0952 | 1.1113 | 1.0808 | 4.5687
B42 | 5756.268 | 1562.976 | 374.969 | 26.8207 | 29.1264 | 27.9496 | 27.9655 | 1.1269 | 1.0851 | 1.1122 | 1.1106 | 1.1088 | 4.5399
B43 | 5677.446 | 1549.716 | 374.187 | 26.4534 | 28.8793 | 27.8913 | 27.7413 | 1.0208 | 0.97873 | 0.98318 | 1.0123 | 0.99892 | 4.9988
B44 | 5835.234 | 1575.932 | 378.972 | 27.1886 | 29.3678 | 28.248 | 28.2681 | 0.9288 | 0.93042 | 0.94197 | 1.0067 | 0.9525 | 5.342
B45 | 5712.483 | 1548.396 | 377.058 | 26.6166 | 28.8547 | 28.1053 | 27.8589 | 1.1153 | 1.0706 | 1.1187 | 1.1267 | 1.108 | 4.5256
B46 | 5741.775 | 1540.62 | 372.06 | 26.7531 | 28.7098 | 27.7328 | 27.7319 | 0.97966 | 1.0008 | 1.0205 | 0.96337 | 0.99134 | 5.0353
B47 | 5751.47 | 1542.3 | 380.24 | 26.7983 | 28.7411 | 28.3425 | 27.9606 | 1.0238 | 1.041 | 0.97803 | 1.0209 | 1.0162 | 4.9527
B48 | 5766.08 | 1538.176 | 375.088 | 26.8664 | 28.6642 | 27.9585 | 27.8297 | 0.92761 | 0.90675 | 0.95281 | 0.9459 | 0.93344 | 5.3665
B49 | 5688.16 | 1516.32 | 373.051 | 26.5033 | 28.2569 | 27.8066 | 27.5223 | 0.96332 | 0.99252 | 0.97035 | 1.0108 | 0.98442 | 5.0324
B50 | 5776.794 | 1545.008 | 377.28 | 26.9163 | 28.7915 | 28.1219 | 27.9432 | 0.93454 | 0.91056 | 0.91745 | 0.95026 | 0.92833 | 5.4181
B51 | 5680.368 | 1516.518 | 373.051 | 26.467 | 28.2606 | 27.8066 | 27.5114 | 0.94016 | 0.92022 | 0.95912 | 0.98182 | 0.9506 | 5.2094
B52 | 5678.424 | 1516.32 | 376.941 | 26.4579 | 28.2569 | 28.0966 | 27.6038 | 0.99643 | 0.96188 | 1.021 | 0.97233 | 0.98818 | 5.0281
B53 | 5712.483 | 1516.907 | 377.315 | 26.6166 | 28.2679 | 28.1245 | 27.6697 | 0.99674 | 0.99101 | 0.98115 | 1.0029 | 0.99298 | 5.0157
B54 | 5721.276 | 1518.853 | 379.86 | 26.6576 | 28.3041 | 28.3142 | 27.7586 | 0.92052 | 0.9093 | 0.94904 | 0.91274 | 0.92303 | 5.4132
B55 | 5682.316 | 1494.116 | 374.229 | 26.4761 | 27.8431 | 27.8944 | 27.4046 | 1.0205 | 1.0154 | 1.0543 | 1.0582 | 1.0373 | 4.7556
B56 | 5819.65 | 1524.31 | 380.952 | 27.116 | 28.4058 | 28.3956 | 27.9725 | 0.92939 | 0.89702 | 0.92624 | 0.91405 | 0.91676 | 5.4922
B57 | 5832.312 | 1529.392 | 382.8 | 27.175 | 28.5005 | 28.5333 | 28.0696 | 0.95777 | 0.96793 | 0.93598 | 0.96711 | 0.95728 | 5.278
B58 | 5935.8 | 1548.043 | 388.8 | 27.6572 | 28.8481 | 28.9805 | 28.4953 | 0.97442 | 0.97481 | 0.96748 | 0.96073 | 0.96938 | 5.2912
B59 | 6005.684 | 1563.822 | 391.79 | 27.9828 | 29.1421 | 29.2034 | 28.7761 | 0.91149 | 0.88105 | 0.85816 | 0.88314 | 0.88366 | 5.8616
B60 | 6129.28 | 1567.166 | 390.861 | 28.5587 | 29.2044 | 29.1342 | 28.9658 | 1.5779 | 1.463 | 1.5289 | 1.5248 | 1.5242 | 3.4208
B61 | 6345.952 | 1586.325 | 395.094 | 29.5682 | 29.5615 | 29.4497 | 29.5265 | 0.89343 | 0.8648 | 0.8944 | 0.90064 | 0.88843 | 5.9822
B62 | 6771.587 | 1699.98 | 415.86 | 31.5514 | 31.6795 | 30.9976 | 31.4095 | 0.88256 | 0.89088 | 0.91758 | 0.91972 | 0.90283 | 6.2622
B63 | 6851.042 | 1773.948 | 444.865 | 31.9216 | 33.0579 | 33.1595 | 32.713 | 0.88521 | 0.87597 | 0.84928 | 0.90823 | 0.87993 | 6.6919
B64 | 5998.48 | 1623.512 | 422.048 | 27.9492 | 30.2545 | 31.4588 | 29.8875 | 0.85923 | 0.8465 | 0.8607 | 0.87729 | 0.861 | 6.2483
C1 | 2302.671 | 557.192 | 142.99 | 10.729 | 10.3834 | 10.6583 | 10.5902 | 1.0815 | 1.1731 | 0.87049 | 0.83238 | 0.99956 | 1.9071
C2 | 2664.366 | 648.375 | 164.964 | 12.4143 | 12.0826 | 12.2962 | 12.2644 | 1.0933 | 1.1017 | 1.048 | 1.0909 | 1.0837 | 2.0372
C3 | 54.999 | 4.995 | 0.996 | 0.25626 | 0.093083 | 0.07424 | 0.14119 | 0.70924 | 0.75446 | 0.64054 | 0.66519 | 0.69372 | 0.036636
C4 | 49.029 | 8 | 0 | 0.22844 | 0.14908 | 0 | 0.12584 | 0.68863 | 0.69398 | 0.59644 | 0.64164 | 0.65636 | 0.034511
C5 | 6084.104 | 1534.866 | 385.833 | 28.3482 | 28.6025 | 28.7594 | 28.57 | 0.89754 | 0.89404 | 0.82389 | 0.85862 | 0.86904 | 5.9176
C6 | 5864.754 | 1502.214 | 378.099 | 27.3261 | 27.994 | 28.1829 | 27.8344 | 0.77021 | 0.78818 | 0.71997 | 0.76217 | 0.76055 | 6.5876
C7 | 5986.292 | 1524.216 | 382.968 | 27.8924 | 28.4041 | 28.5458 | 28.2808 | 0.75642 | 0.80742 | 0.71565 | 0.73758 | 0.75503 | 6.7421
C8 | 5846.988 | 1505.175 | 374.957 | 27.2433 | 28.0492 | 27.9487 | 27.7471 | 0.83728 | 0.84456 | 0.79177 | 0.77644 | 0.81303 | 6.143
C9 | 59.02 | 3.99 | 2 | 0.275 | 0.074354 | 0.14908 | 0.16614 | 0.77135 | 0.76463 | 0.69522 | 0.6952 | 0.73251 | 0.040826
C10 | 51.975 | 7.005 | 0.994 | 0.24217 | 0.13054 | 0.074091 | 0.14893 | 0.80479 | 0.79897 | 0.73413 | 0.78863 | 0.78213 | 0.034276
C11 | 5973.188 | 1532.902 | 378.3 | 27.8314 | 28.5659 | 28.1979 | 28.1984 | 0.83833 | 0.84214 | 0.82089 | 0.82319 | 0.83119 | 6.1066
C12 | 5812.485 | 1535.202 | 378.3 | 27.0826 | 28.6088 | 28.1979 | 27.9631 | 0.80374 | 0.79044 | 0.73691 | 0.79814 | 0.78276 | 6.4302
C13 | 5911.872 | 1527.992 | 371.712 | 27.5457 | 28.4744 | 27.7068 | 27.909 | 0.857 | 0.88302 | 0.79707 | 0.84535 | 0.84618 | 5.9368
C14 | 5814.455 | 1528.3 | 375.192 | 27.0918 | 28.4802 | 27.9662 | 27.8461 | 0.81276 | 0.79525 | 0.77447 | 0.79656 | 0.79488 | 6.3058
C15 | 6036.08 | 1556.24 | 381.21 | 28.1244 | 29.0008 | 28.4148 | 28.5133 | 0.81216 | 0.84829 | 0.76645 | 0.80788 | 0.80922 | 6.3424
C16 | 5833.122 | 1539.58 | 379.08 | 27.1787 | 28.6904 | 28.256 | 28.0417 | 0.79622 | 0.79774 | 0.76379 | 0.84115 | 0.8002 | 6.3078
C17 | 6076.616 | 1568.358 | 385.704 | 28.3133 | 29.2267 | 28.7498 | 28.7632 | 0.77431 | 0.77797 | 0.72305 | 0.76375 | 0.76008 | 6.8116
C18 | 44 | 11.001 | 2.002 | 0.20501 | 0.20501 | 0.14923 | 0.18641 | 0.65291 | 0.69033 | 0.61291 | 0.67979 | 0.65966 | 0.050867
C19 | 45.98 | 8.993 | 2.004 | 0.21424 | 0.16759 | 0.14938 | 0.17707 | 0.75104 | 0.76529 | 0.72697 | 0.7975 | 0.76063 | 0.041902
C20 | 5826.206 | 1548.999 | 381.225 | 27.1465 | 28.8659 | 28.4159 | 28.1428 | 0.83085 | 0.83725 | 0.79865 | 0.85235 | 0.83001 | 6.1032
C21 | 5972.328 | 1545.46 | 381.22 | 27.8274 | 28.7999 | 28.4155 | 28.3476 | 0.83216 | 0.83529 | 0.81009 | 0.89015 | 0.84244 | 6.0569
C22 | 5687.92 | 1513.12 | 372.681 | 26.5022 | 28.1973 | 27.7791 | 27.4928 | 0.78412 | 0.79928 | 0.76786 | 0.8639 | 0.80462 | 6.1504
C23 | 5921.316 | 1535.66 | 378.3 | 27.5897 | 28.6173 | 28.1979 | 28.135 | 0.86955 | 0.86257 | 0.84799 | 0.92727 | 0.87736 | 5.7722
C24 | 5698.7 | 1517.04 | 374.99 | 26.5524 | 28.2703 | 27.9512 | 27.5913 | 0.99697 | 0.96939 | 0.97067 | 1.0729 | 1.0034 | 4.9497
C25 | 5964.48 | 1546.218 | 381.03 | 27.7908 | 28.8141 | 28.4014 | 28.3354 | 0.89807 | 0.91992 | 0.90112 | 1.0029 | 0.93148 | 5.4756
C26 | 5580.3 | 1496.764 | 368.046 | 26.0007 | 27.8925 | 27.4336 | 27.1089 | 0.91723 | 0.89612 | 0.85864 | 0.96173 | 0.90919 | 5.367
C27 | 6012.3 | 1551.225 | 378.886 | 28.0136 | 28.9074 | 28.2416 | 28.3875 | 0.95962 | 0.94929 | 0.95725 | 1.0302 | 0.97465 | 5.2427
C28 | 5699.68 | 1537.798 | 379.076 | 26.557 | 28.6572 | 28.2557 | 27.8233 | 0.91963 | 0.84507 | 0.92619 | 0.99939 | 0.92419 | 5.419
C29 | 5941.917 | 1545.124 | 378.852 | 27.6857 | 28.7937 | 28.239 | 28.2395 | 1.127 | 1.1989 | 1.0712 | 1.1301 | 1.1327 | 4.4875
C30 | 5640.126 | 1540.128 | 379.08 | 26.2795 | 28.7006 | 28.256 | 27.7454 | 1.0442 | 1.0422 | 1.0087 | 1.1323 | 1.0578 | 4.7211
C31 | 5939.78 | 1534.35 | 370.048 | 27.6757 | 28.5929 | 27.5828 | 27.9505 | 2.1231 | 1.4578 | 2.0893 | 2.5645 | 2.0961 | 2.4002
C32 | 5291.495 | 1501.23 | 365.956 | 24.6551 | 27.9757 | 27.2778 | 26.6362 | 1.1642 | 1.1977 | 0.99596 | 1.157 | 1.1314 | 4.2376
C33 | 0.999 | 1 | 1.998 | 0.0046547 | 0.018635 | 0.14893 | 0.057406 | 0.68165 | 0.68463 | 0.70066 | 0.68404 | 0.68779 | 0.015024
C34 | 2 | 3.003 | 1 | 0.0093188 | 0.055961 | 0.074538 | 0.046606 | 0.67855 | 0.6547 | 0.67735 | 0.67072 | 0.6704 | 0.012514
C35 | 2.002 | 1 | 1.001 | 0.0093281 | 0.018635 | 0.074613 | 0.034192 | 0.65648 | 0.62941 | 0.65061 | 0.65642 | 0.64833 | 0.009493
C36 | 2 | 1.001 | 0 | 0.0093188 | 0.018654 | 0 | 0.0093242 | 0.73117 | 0.72154 | 0.73791 | 0.74117 | 0.73299 | 0.0022897
C37 | 1 | 1.001 | 1 | 0.0046594 | 0.018654 | 0.074538 | 0.032617 | 0.62465 | 0.6376 | 0.64561 | 0.63633 | 0.63609 | 0.00923
C38 | 0.996 | 0 | 0 | 0.0046407 | 0 | 0 | 0.0015469 | 0.61447 | 0.6291 | 0.62072 | 0.60594 | 0.61761 | 0.00045084
C39 | 1 | 2.002 | 0 | 0.0046594 | 0.037308 | 0 | 0.013989 | 0.68347 | 0.65976 | 0.68835 | 0.66968 | 0.67541 | 0.0037281
C40 | 0.999 | 0 | 0 | 0.0046547 | 0 | 0 | 0.0015516 | 0.63808 | 0.65844 | 0.63592 | 0.64264 | 0.64383 | 0.00043378
C41 | 1.001 | 2.002 | 0 | 0.004664 | 0.037308 | 0 | 0.013991 | 0.65924 | 0.7002 | 0.65654 | 0.68588 | 0.67572 | 0.0037269
C42 | 2 | 0 | 1 | 0.0093188 | 0 | 0.074538 | 0.027952 | 0.57366 | 0.57715 | 0.59308 | 0.58133 | 0.58135 | 0.0086547
C43 | 1.001 | 0.999 | 1.001 | 0.004664 | 0.018617 | 0.074613 | 0.032631 | 0.60524 | 0.64897 | 0.65373 | 0.63666 | 0.63643 | 0.009229
C44 | 2.996 | 1.001 | 1 | 0.01396 | 0.018654 | 0.074538 | 0.035717 | 0.63009 | 0.62403 | 0.60167 | 0.61288 | 0.61726 | 0.010416
C45 | 0 | 1.001 | 3 | 0 | 0.018654 | 0.22362 | 0.080756 | 0.64034 | 0.63266 | 0.63924 | 0.64557 | 0.63947 | 0.022732
C46 | 4.008 | 2.002 | 2.002 | 0.018675 | 0.037308 | 0.14923 | 0.068403 | 0.63379 | 0.69157 | 0.66769 | 0.65458 | 0.66224 | 0.018592
C47 | 3.003 | 0 | 0 | 0.013992 | 0 | 0 | 0.004664 | 0.69662 | 0.67152 | 0.70138 | 0.69435 | 0.69106 | 0.0012148
C48 | 2.002 | 0 | 0 | 0.0093281 | 0 | 0 | 0.0031094 | 0.67931 | 0.66488 | 0.65467 | 0.66749 | 0.66665 | 0.00083955
C49 | 2.002 | 1.998 | 0 | 0.0093281 | 0.037233 | 0 | 0.01552 | 0.66599 | 0.66226 | 0.68325 | 0.686 | 0.67446 | 0.0041421
C50 | 6 | 0 | 1 | 0.027956 | 0 | 0.074538 | 0.034165 | 0.59033 | 0.62455 | 0.61529 | 0.61634 | 0.61176 | 0.010052
C51 | 3 | 1 | 1 | 0.013978 | 0.018635 | 0.074538 | 0.035717 | 0.58292 | 0.57131 | 0.57881 | 0.57216 | 0.57632 | 0.011155
C52 | 4.998 | 0 | 1.001 | 0.023288 | 0 | 0.074613 | 0.032634 | 0.67471 | 0.68014 | 0.68074 | 0.68055 | 0.67904 | 0.0086505
C53 | 4.995 | 2.002 | 1.001 | 0.023274 | 0.037308 | 0.074613 | 0.045065 | 0.73003 | 0.7564 | 0.7272 | 0.734 | 0.737 | 0.011006
C54 | 6 | 0.996 | 1.998 | 0.027956 | 0.018561 | 0.14893 | 0.065148 | 0.69197 | 0.71022 | 0.71564 | 0.68599 | 0.70106 | 0.016727
C55 | 5.005 | 1.001 | 1 | 0.02332 | 0.018654 | 0.074538 | 0.038837 | 0.65245 | 0.65175 | 0.63286 | 0.62907 | 0.64162 | 0.010896
C56 | 7.99 | 3.003 | 1.001 | 0.037228 | 0.055961 | 0.074613 | 0.055934 | 0.77905 | 0.78719 | 0.78238 | 0.74104 | 0.77264 | 0.013031
C57 | 6.006 | 0.999 | 1.001 | 0.027984 | 0.018617 | 0.074613 | 0.040405 | 0.71747 | 0.70381 | 0.696 | 0.67834 | 0.69905 | 0.010404
C58 | 8.987 | 1.998 | 0.999 | 0.041874 | 0.037233 | 0.074464 | 0.05119 | 0.65199 | 0.68105 | 0.66996 | 0.66779 | 0.66778 | 0.013798
C59 | 7 | 0 | 1.005 | 0.032616 | 0 | 0.074911 | 0.035842 | 0.69902 | 0.70673 | 0.70229 | 0.70853 | 0.70415 | 0.0091623
C60 | 8 | 3.003 | 0 | 0.037275 | 0.055961 | 0 | 0.031079 | 0.69114 | 0.68325 | 0.6896 | 0.66799 | 0.68306 | 0.00819
C61 | 11 | 0.999 | 0.999 | 0.051253 | 0.018617 | 0.074464 | 0.048111 | 0.73139 | 0.77184 | 0.73223 | 0.72396 | 0.74009 | 0.011701
C62 | 10 | 2.01 | 2 | 0.046594 | 0.037457 | 0.14908 | 0.077709 | 0.70322 | 0.73811 | 0.70923 | 0.70777 | 0.71472 | 0.019571
C63 | 10.983 | 2.002 | 1 | 0.051174 | 0.037308 | 0.074538 | 0.05434 | 0.64655 | 0.66365 | 0.63918 | 0.64108 | 0.64769 | 0.015102
C64 | 10.005 | 4.004 | 2.002 | 0.046617 | 0.074615 | 0.14923 | 0.090153 | 0.69338 | 0.72066 | 0.70868 | 0.68164 | 0.70125 | 0.023141
D1 | 2221.294 | 531.92 | 134.028 | 10.3499 | 9.9124 | 9.9902 | 10.0842 | 0.72442 | 0.78067 | 0.62192 | 0.6405 | 0.69485 | 2.6123
D2 | 2532.005 | 616.341 | 152.886 | 11.7976 | 11.4856 | 11.3959 | 11.5597 | 0.79741 | 0.77138 | 0.82401 | 0.80763 | 0.80034 | 2.5998
D3 | 57.035 | 2.002 | 1.001 | 0.26575 | 0.037308 | 0.074613 | 0.12589 | 0.62185 | 0.6351 | 0.63408 | 0.66711 | 0.63975 | 0.03542
D4 | 54.993 | 3.009 | 2.006 | 0.25623 | 0.056073 | 0.14952 | 0.15394 | 0.80886 | 0.8725 | 0.71431 | 0.70668 | 0.77864 | 0.035587
D5 | 5843.032 | 1466.665 | 369.904 | 27.2249 | 27.3316 | 27.5721 | 27.3762 | 0.81556 | 0.80319 | 0.83467 | 0.84174 | 0.82393 | 5.9808
D6 | 5725.17 | 1474.599 | 369.904 | 26.6758 | 27.4794 | 27.5721 | 27.2424 | 0.74474 | 0.74475 | 0.74231 | 0.74934 | 0.74529 | 6.5795
D7 | 5775.76 | 1466.665 | 368.172 | 26.9115 | 27.3316 | 27.443 | 27.2287 | 0.78004 | 0.76835 | 0.84513 | 0.86641 | 0.81605 | 6.006
D8 | 5785.458 | 1490.493 | 369.012 | 26.9567 | 27.7756 | 27.5056 | 27.4126 | 0.87978 | 0.90732 | 0.92234 | 0.93275 | 0.91076 | 5.4177
D9 | 65.959 | 1.001 | 3.006 | 0.30733 | 0.018654 | 0.22406 | 0.18335 | 0.6932 | 0.73921 | 0.69001 | 0.6753 | 0.69984 | 0.047158
D10 | 57.024 | 1 | 1 | 0.2657 | 0.018635 | 0.074538 | 0.11962 | 0.65013 | 0.67124 | 0.55709 | 0.58655 | 0.61798 | 0.034843
D11 | 5767.041 | 1483.41 | 364.219 | 26.8708 | 27.6436 | 27.1483 | 27.2209 | 0.83919 | 0.86299 | 0.89669 | 0.90678 | 0.87682 | 5.5881
D12 | 4683.995 | 1477.539 | 366.976 | 21.8245 | 27.5342 | 27.3538 | 25.5709 | 0.97201 | 1.0003 | 0.96229 | 0.98377 | 0.97969 | 4.6982
D13 | 5759.052 | 1483.461 | 364.125 | 26.8336 | 27.6446 | 27.1413 | 27.2065 | 0.78719 | 0.78619 | 0.80113 | 0.82898 | 0.80106 | 6.1134
D14 | 5647.408 | 1477.449 | 363.805 | 26.3134 | 27.5325 | 27.1175 | 26.9878 | 0.83342 | 0.8758 | 0.83984 | 0.90145 | 0.86307 | 5.6285
D15 | 5764.08 | 1501.227 | 366.872 | 26.857 | 27.9757 | 27.3461 | 27.3929 | 0.72861 | 0.76442 | 0.7791 | 0.84278 | 0.77982 | 6.3229
D16 | 5794.551 | 1517.126 | 369.951 | 26.999 | 28.2719 | 27.5756 | 27.6155 | 0.74832 | 0.78681 | 0.73973 | 0.80288 | 0.76988 | 6.4566
D17 | 5770.908 | 1510.005 | 365.956 | 26.8889 | 28.1392 | 27.2778 | 27.4353 | 0.82211 | 0.86616 | 0.80509 | 0.85037 | 0.83627 | 5.9052
D18 | 47.023 | 2.002 | 3.003 | 0.2191 | 0.037308 | 0.22384 | 0.16008 | 0.6894 | 0.70884 | 0.61856 | 0.65639 | 0.66918 | 0.04306
D19 | 52.038 | 0 | 1.005 | 0.24246 | 0 | 0.074911 | 0.10579 | 0.79649 | 0.83266 | 0.77782 | 0.77769 | 0.79648 | 0.023908
D20 | 5902.312 | 1561.04 | 383.899 | 27.5011 | 29.0903 | 28.6152 | 28.4022 | 0.97347 | 0.98564 | 0.91365 | 0.98267 | 0.9643 | 5.3016
D21 | 5978.118 | 1565.919 | 386 | 27.8543 | 29.1812 | 28.7718 | 28.6025 | 0.82141 | 0.83027 | 0.75919 | 0.76503 | 0.79463 | 6.4791
D22 | 5835.144 | 1545.232 | 379.064 | 27.1882 | 28.7957 | 28.2548 | 28.0796 | 0.77582 | 0.83509 | 0.76162 | 0.76546 | 0.78506 | 6.4381
D23 | 5935.818 | 1556.688 | 380.16 | 27.6572 | 29.0092 | 28.3365 | 28.3343 | 0.84256 | 0.91522 | 0.8759 | 0.84596 | 0.8704 | 5.8596
D24 | 5740.156 | 1520.84 | 373.842 | 26.7456 | 28.3411 | 27.8656 | 27.6508 | 0.85994 | 0.91385 | 0.80814 | 0.82483 | 0.85265 | 5.8373
D25 | 5827.248 | 1527.992 | 375.751 | 27.1514 | 28.4744 | 28.0079 | 27.8779 | 0.87346 | 0.92181 | 0.77507 | 0.82095 | 0.84961 | 5.9063
D26 | 5670.132 | 1507.149 | 371.856 | 26.4193 | 28.086 | 27.7176 | 27.4076 | 0.75682 | 0.80034 | 0.69246 | 0.73308 | 0.7467 | 6.6069
D27 | 5896.338 | 1549.8 | 381.996 | 27.4733 | 28.8808 | 28.4734 | 28.2758 | 0.83181 | 0.84374 | 0.76128 | 0.78283 | 0.80563 | 6.3176
D28 | 5835.128 | 1549.992 | 380.952 | 27.1881 | 28.8844 | 28.3956 | 28.156 | 0.80195 | 0.85331 | 0.76878 | 0.72795 | 0.78933 | 6.4207
D29 | 5918.052 | 1562.733 | 385.884 | 27.5745 | 29.1218 | 28.7632 | 28.4865 | 0.92445 | 0.94171 | 0.9117 | 0.95834 | 0.93422 | 5.4886
D30 | 5773.95 | 1553.936 | 382.876 | 26.903 | 28.9579 | 28.539 | 28.1333 | 1.1013 | 1.1575 | 0.90249 | 0.92061 | 1.0265 | 4.9333
D31 | 5966.286 | 1575.384 | 387.283 | 27.7992 | 29.3576 | 28.8675 | 28.6748 | 0.80718 | 0.87043 | 0.74645 | 0.81748 | 0.81158 | 6.3597
D32 | 5817.378 | 1561.225 | 384.938 | 27.1054 | 29.0937 | 28.6927 | 28.2973 | 0.86351 | 0.88046 | 0.77456 | 0.79494 | 0.82957 | 6.14
D33 | 5914.22 | 1562.21 | 383.961 | 27.5566 | 29.1121 | 28.6199 | 28.4295 | 0.78696 | 0.7688 | 0.76492 | 0.94011 | 0.81842 | 6.2526
D34 | 5704.86 | 1564.782 | 380.78 | 26.5811 | 29.16 | 28.3827 | 28.0413 | 0.90142 | 0.91467 | 0.89408 | 1.0316 | 0.93712 | 5.3861
D35 | 5911.143 | 1567.135 | 385.09 | 27.5423 | 29.2039 | 28.704 | 28.4834 | 0.81918 | 0.80474 | 0.82285 | 0.93053 | 0.84581 | 6.0616
D36 | 29.988 | 1.992 | 0 | 0.13973 | 0.037121 | 0 | 0.058949 | 0.65949 | 0.66295 | 0.67706 | 0.68252 | 0.67057 | 0.015824
D37 | 24 | 1.001 | 1 | 0.11183 | 0.018654 | 0.074538 | 0.068339 | 0.61524 | 0.65524 | 0.6205 | 0.60948 | 0.62537 | 0.01967
D38 | 5761.028 | 1549.405 | 381.808 | 26.8428 | 28.8735 | 28.4594 | 28.0586 | 0.75454 | 0.73818 | 0.69622 | 0.78414 | 0.74395 | 6.7888
D39 | 5888.48 | 1562.81 | 384.747 | 27.4367 | 29.1233 | 28.6784 | 28.4128 | 0.84039 | 0.88129 | 0.82171 | 0.87739 | 0.85556 | 5.9777
D40 | 5768.932 | 1554.922 | 383.961 | 26.8797 | 28.9763 | 28.6199 | 28.1586 | 0.85241 | 0.86297 | 0.82516 | 0.9141 | 0.86426 | 5.8646
D41 | 5921.013 | 1569.932 | 386.631 | 27.5883 | 29.256 | 28.8189 | 28.5544 | 0.89093 | 0.87272 | 0.86631 | 0.92256 | 0.8884 | 5.7855
D42 | 5841.056 | 1569.48 | 390.173 | 27.2157 | 29.2476 | 29.0829 | 28.5154 | 0.84589 | 0.82028 | 0.79773 | 0.82813 | 0.82319 | 6.2352
D43 | 5892.39 | 1563.195 | 387.828 | 27.4549 | 29.1304 | 28.9081 | 28.4978 | 0.77531 | 0.82572 | 0.7699 | 0.80048 | 0.79317 | 6.4673
D44 | 5812.443 | 1561.04 | 387.869 | 27.0824 | 29.0903 | 28.9112 | 28.3613 | 0.76344 | 0.77846 | 0.74884 | 0.77797 | 0.76727 | 6.6535
D45 | 5942.901 | 1573.884 | 390.015 | 27.6902 | 29.3296 | 29.0711 | 28.697 | 0.81582 | 0.79436 | 0.78499 | 0.85687 | 0.81348 | 6.3498
D46 | 5844.96 | 1559.038 | 391.88 | 27.2339 | 29.053 | 29.2101 | 28.499 | 0.81204 | 0.86413 | 0.79669 | 0.85812 | 0.83325 | 6.1564
D47 | 5818.161 | 1538.316 | 382.641 | 27.109 | 28.6668 | 28.5215 | 28.0991 | 0.73859 | 0.75781 | 0.71996 | 0.73964 | 0.73912 | 6.8431
D48 | 5773.566 | 1538.976 | 386.163 | 26.9012 | 28.6791 | 28.784 | 28.1215 | 0.92583 | 0.89415 | 0.94563 | 0.94318 | 0.92743 | 5.458
D49 | 5747.079 | 1511.097 | 380.25 | 26.7778 | 28.1596 | 28.3432 | 27.7602 | 0.8401 | 0.8657 | 0.84229 | 0.86571 | 0.85354 | 5.8543
D50 | 5390.05 | 1498.266 | 377.33 | 25.1143 | 27.9205 | 28.1256 | 27.0535 | 0.77964 | 0.7897 | 0.77143 | 0.80307 | 0.78605 | 6.1951
D51 | 5780.788 | 1510.005 | 381.03 | 26.9349 | 28.1392 | 28.4014 | 27.8252 | 0.74991 | 0.76191 | 0.72963 | 0.73252 | 0.74361 | 6.7355
D52 | 5596.032 | 1525.038 | 383.94 | 26.074 | 28.4194 | 28.6183 | 27.7039 | 0.78845 | 0.76666 | 0.75791 | 0.76506 | 0.7696 | 6.4796
D53 | 309.872 | 76.006 | 11.024 | 1.4438 | 1.4164 | 0.82171 | 1.2273 | 2.584 | 2.5539 | 2.8556 | 2.9715 | 2.747 | 0.080421
D54 | 5581.62 | 1470.63 | 370.922 | 26.0069 | 27.4055 | 27.6479 | 27.0201 | 0.76631 | 0.79249 | 0.80649 | 0.82443 | 0.79771 | 6.097
D55 | 5757.84 | 1498.335 | 376.938 | 26.828 | 27.9218 | 28.0964 | 27.6154 | 0.76181 | 0.79805 | 0.75135 | 0.77154 | 0.77088 | 6.4482
D56 | 5532.753 | 1455.808 | 367.038 | 25.7792 | 27.1293 | 27.3584 | 26.7556 | 0.81858 | 0.80606 | 0.8503 | 0.81311 | 0.82219 | 5.8576
D57 | 4892.583 | 1274.898 | 322.31 | 22.7964 | 23.758 | 24.0245 | 23.5263 | 0.83867 | 0.83584 | 0.81561 | 0.82222 | 0.82814 | 5.1135
D58 | 4738.962 | 1245.151 | 314.916 | 22.0806 | 23.2036 | 23.4733 | 22.9192 | 0.78115 | 0.76887 | 0.80698 | 0.80547 | 0.79079 | 5.2169
D59 | 5689.53 | 1479.513 | 374.016 | 26.5097 | 27.571 | 27.8786 | 27.3198 | 0.73058 | 0.74149 | 0.7397 | 0.73362 | 0.73636 | 6.6782
D60 | 5670.502 | 1485 | 374.22 | 26.421 | 27.6733 | 27.8938 | 27.3294 | 0.86768 | 0.85423 | 0.84887 | 0.87572 | 0.86169 | 5.7089
D61 | 6091.872 | 1537.746 | 386.12 | 28.3844 | 28.6562 | 28.7808 | 28.6071 | 0.86445 | 0.87892 | 0.85563 | 0.9162 | 0.8791 | 5.8574
D62 | 6273.63 | 1588.446 | 393.092 | 29.2312 | 29.601 | 29.3005 | 29.3776 | 0.82547 | 0.86803 | 0.81431 | 0.86694 | 0.84403 | 6.2651
D63 | 6851.7 | 1713.937 | 423.907 | 31.9247 | 31.9395 | 31.5974 | 31.8205 | 0.80493 | 0.80548 | 0.78748 | 0.81336 | 0.80287 | 7.134
D64 | 4631.487 | 1271.94 | 322.31 | 21.5799 | 23.7028 | 24.0245 | 23.1024 | 0.86926 | 0.84251 | 0.8559 | 0.80823 | 0.84428 | 4.9254
E1 | 2193.602 | 532.885 | 134.9 | 10.2208 | 9.9304 | 10.0552 | 10.0688 | 0.70621 | 0.72341 | 0.72099 | 0.72902 | 0.71996 | 2.5174
E2 | 2297.447 | 567.45 | 145.068 | 10.7047 | 10.5745 | 10.8131 | 10.6975 | 1.2007 | 1.2544 | 1.2299 | 1.2278 | 1.2283 | 1.5676
E3 | 59.027 | 1 | 0 | 0.27503 | 0.018635 | 0 | 0.097888 | 0.63694 | 0.65959 | 0.62483 | 0.65362 | 0.64389 | 0.027365
E4 | 65.016 | 4.004 | 0 | 0.30293 | 0.074615 | 0 | 0.12585 | 0.70655 | 0.7139 | 0.69377 | 0.70676 | 0.70528 | 0.032119
E5 | 6017.076 | 1521.954 | 385.704 | 28.0359 | 28.3619 | 28.7498 | 28.3825 | 0.85425 | 0.85753 | 0.83017 | 0.85897 | 0.85031 | 6.0082
E6 | 5982.461 | 1509.02 | 381.81 | 27.8746 | 28.1209 | 28.4595 | 28.1517 | 0.72845 | 0.73624 | 0.73678 | 0.76518 | 0.7418 | 6.8311
E7 | 5913.117 | 1501.227 | 378.099 | 27.5515 | 27.9757 | 28.1829 | 27.9033 | 0.71266 | 0.70498 | 0.70668 | 0.6868 | 0.70285 | 7.1461
E8 | 5929.055 | 1509.214 | 376.938 | 27.6257 | 28.1245 | 28.0964 | 27.9489 | 0.80155 | 0.78258 | 0.79153 | 0.79464 | 0.79261 | 6.3472
E9 | 74 | 3 | 0 | 0.34479 | 0.055906 | 0 | 0.13357 | 0.59632 | 0.59329 | 0.59582 | 0.58853 | 0.5935 | 0.040509
E10 | 78.948 | 5.005 | 1.001 | 0.36785 | 0.093269 | 0.074613 | 0.17858 | 0.56245 | 0.58155 | 0.5561 | 0.56451 | 0.56623 | 0.056768
E11 | 5778.885 | 1490.37 | 369.146 | 26.926 | 27.7733 | 27.5156 | 27.405 | 0.82165 | 0.8244 | 0.8158 | 0.78793 | 0.81257 | 6.0707
E12 | 5839.092 | 1502.214 | 369.951 | 27.2066 | 27.994 | 27.5756 | 27.5921 | 0.74757 | 0.73175 | 0.77698 | 0.76976 | 0.75673 | 6.5632
E13 | 5778.727 | 1499.253 | 368.104 | 26.9253 | 27.9389 | 27.4379 | 27.434 | 0.73703 | 0.77044 | 0.7885 | 0.76225 | 0.76478 | 6.4569
E14 | 5698.44 | 1471.59 | 361.956 | 26.5512 | 27.4234 | 26.9796 | 26.9847 | 0.75578 | 0.75501 | 0.71227 | 0.74405 | 0.74199 | 6.5463
E15 | 5682.976 | 1484.489 | 360.983 | 26.4792 | 27.6637 | 26.9071 | 27.0167 | 0.78589 | 0.78995 | 0.79605 | 0.81553 | 0.79694 | 6.1021
E16 | 5595.186 | 1452.384 | 353.925 | 26.0701 | 27.0655 | 26.381 | 26.5055 | 0.77688 | 0.83332 | 0.8006 | 0.81612 | 0.807 | 5.912
E17 | 5640.03 | 1486.365 | 362.838 | 26.2791 | 27.6987 | 27.0454 | 27.0077 | 0.7607 | 0.75598 | 0.71569 | 0.76911 | 0.75065 | 6.4762
E18 | 100.062 | 5.992 | 1.001 | 0.46623 | 0.11166 | 0.074613 | 0.2175 | 0.79901 | 0.82356 | 0.80921 | 0.81915 | 0.81279 | 0.048168
E19 | 101.043 | 6.006 | 1 | 0.4708 | 0.11192 | 0.074538 | 0.21909 | 0.62221 | 0.65679 | 0.62232 | 0.65619 | 0.63961 | 0.061656
E20 | 5528.187 | 1452.864 | 355.264 | 25.7579 | 27.0744 | 26.4808 | 26.4377 | 0.76966 | 0.72376 | 0.75313 | 0.76122 | 0.75214 | 6.327
E21 | 5719.387 | 1513.071 | 370.214 | 26.6488 | 28.1964 | 27.5952 | 27.4801 | 0.6891 | 0.69803 | 0.70796 | 0.69448 | 0.69743 | 7.0924
E22 | 5799.612 | 1496.664 | 367.198 | 27.0226 | 27.8906 | 27.3704 | 27.4279 | 0.70031 | 0.7239 | 0.67528 | 0.71393 | 0.70359 | 7.0169
E23 | 4115.088 | 1513.945 | 376.915 | 19.1738 | 28.2127 | 28.0947 | 25.1604 | 0.76899 | 0.7328 | 0.75871 | 0.75649 | 0.75436 | 6.0036
E24 | 5895.294 | 1549.8 | 380.052 | 27.4684 | 28.8808 | 28.3285 | 28.2259 | 0.72255 | 0.70522 | 0.70281 | 0.75872 | 0.72267 | 7.0304
E25 | 5799.612 | 1556.894 | 383.165 | 27.0226 | 29.013 | 28.5605 | 28.1987 | 0.79461 | 0.83801 | 0.81091 | 0.80984 | 0.81349 | 6.2395
E26 | 5963.454 | 1571.075 | 384.944 | 27.786 | 29.2773 | 28.6931 | 28.5855 | 0.70346 | 0.68098 | 0.66226 | 0.70608 | 0.68842 | 7.4741
E27 | 5889.378 | 1567.356 | 384.938 | 27.4409 | 29.208 | 28.6927 | 28.4472 | 0.71701 | 0.7054 | 0.74652 | 0.75026 | 0.73005 | 7.0139
E28 | 6043.401 | 1592.745 | 392.178 | 28.1585 | 29.6811 | 29.2323 | 29.024 | 0.69611 | 0.74627 | 0.72334 | 0.71208 | 0.71968 | 7.2592
E29 | 5970.593 | 1591.044 | 390.975 | 27.8193 | 29.6494 | 29.1427 | 28.8705 | 0.71051 | 0.72659 | 0.73085 | 0.71749 | 0.7214 | 7.2036
E30 | 5969.496 | 1574.03 | 387.869 | 27.8142 | 29.3324 | 28.9112 | 28.6859 | 0.79212 | 0.79469 | 0.77473 | 0.82119 | 0.79586 | 6.4879
E31 | 5939.934 | 1574.03 | 388.872 | 27.6764 | 29.3324 | 28.9859 | 28.6649 | 0.74673 | 0.70901 | 0.73959 | 0.74293 | 0.73472 | 7.0227
E32 | 5778.946 | 1528.152 | 376.145 | 26.9263 | 28.4774 | 28.0373 | 27.8137 | 0.81064 | 0.79586 | 0.79523 | 0.78852 | 0.79761 | 6.2769
E33 | 5754.002 | 1532.95 | 376.748 | 26.8101 | 28.5668 | 28.0822 | 27.8197 | 0.77277 | 0.76696 | 0.77411 | 0.76806 | 0.77048 | 6.4993
E34 | 5454.748 | 1439.085 | 356.118 | 25.4158 | 26.8176 | 26.5445 | 26.2593 | 0.79691 | 0.76083 | 0.77979 | 0.75649 | 0.77367 | 6.1094
E35 | 5249.464 | 1399.134 | 345.933 | 24.4593 | 26.0731 | 25.7853 | 25.4392 | 0.77587 | 0.77318 | 0.74833 | 0.75864 | 0.76409 | 5.9929
E36 | 25.976 | 1.998 | 0.999 | 0.12103 | 0.037233 | 0.074464 | 0.077576 | 0.7507 | 0.69967 | 0.71439 | 0.72423 | 0.72249 | 0.019327
E37 | 35.01 | 2.002 | 1 | 0.16312 | 0.037308 | 0.074538 | 0.091657 | 0.71341 | 0.71317 | 0.67394 | 0.69528 | 0.69914 | 0.023598
E38 | 4777.25 | 1265.038 | 312.012 | 22.259 | 23.5742 | 23.2569 | 23.03 | 0.69246 | 0.73841 | 0.70919 | 0.70442 | 0.71132 | 5.8278
E39 | 4766.324 | 1276.191 | 318.825 | 22.2081 | 23.7821 | 23.7647 | 23.2516 | 0.7389 | 0.71665 | 0.73258 | 0.77318 | 0.74061 | 5.6511
E40 | 44.997 | 6 | 0.999 | 0.20966 | 0.11181 | 0.074464 | 0.13198 | 0.67484 | 0.66589 | 0.67419 | 0.67889 | 0.67347 | 0.035274
E41 | 4808.77 | 1275.934 | 319.746 | 22.4059 | 23.7773 | 23.8334 | 23.3388 | 0.72444 | 0.72439 | 0.71024 | 0.72882 | 0.722 | 5.8185
E42 | 4871.828 | 1287.048 | 320.739 | 22.6997 | 23.9844 | 23.9074 | 23.5305 | 0.74374 | 0.7548 | 0.72662 | 0.74905 | 0.74363 | 5.6957
E43 | 4939.066 | 1308.422 | 327.227 | 23.013 | 24.3827 | 24.391 | 23.9289 | 0.76306 | 0.74577 | 0.76376 | 0.76034 | 0.75827 | 5.6803
E44 | 61.037 | 8.993 | 0.999 | 0.28439 | 0.16759 | 0.074464 | 0.17548 | 0.64547 | 0.65916 | 0.63759 | 0.6292 | 0.64295 | 0.049128
E45 | 5237.388 | 1392.092 | 350.17 | 24.403 | 25.9419 | 26.1011 | 25.482 | 0.72952 | 0.70955 | 0.7239 | 0.74832 | 0.72795 | 6.3009
E46 | 677.06 | 168.912 | 39.008 | 3.1547 | 3.1477 | 2.9076 | 3.07 | 1.8684 | 1.7755 | 1.6989 | 2.0609 | 1.8559 | 0.29776
E47 | 68.952 | 11 | 1 | 0.32127 | 0.20499 | 0.074538 | 0.20027 | 0.68307 | 0.68013 | 0.67235 | 0.65728 | 0.67328 | 0.053541
E48 | 64.008 | 9.99 | 1.005 | 0.29824 | 0.18617 | 0.074911 | 0.18644 | 0.67149 | 0.65484 | 0.65849 | 0.67715 | 0.66555 | 0.050422
E49 | 5639.504 | 1487.35 | 377.122 | 26.2766 | 27.7171 | 28.1101 | 27.3679 | 0.74966 | 0.77614 | 0.72149 | 0.73877 | 0.74678 | 6.5966
E50 | 108.025 | 14.004 | 1.995 | 0.50333 | 0.26097 | 0.1487 | 0.30433 | 1.7588 | 1.6392 | 1.6001 | 2.0135 | 1.7603 | 0.03112
E51 | 436.08 | 99.968 | 23.994 | 2.0319 | 1.8629 | 1.7885 | 1.8944 | 1.9257 | 1.913 | 1.8326 | 2.1515 | 1.9593 | 0.17404
E52 | 1521.956 | 415.922 | 106.026 | 7.0914 | 7.7508 | 7.903 | 7.5817 | 1.7841 | 1.812 | 1.6219 | 1.8029 | 1.7569 | 0.77676
E53 | 57.063 | 8.993 | 0.996 | 0.26588 | 0.16759 | 0.07424 | 0.16923 | 0.65308 | 0.65975 | 0.67108 | 0.66518 | 0.66231 | 0.045994
E54 | 5355.948 | 1402.96 | 357.192 | 24.9554 | 26.1444 | 26.6245 | 25.9081 | 0.81721 | 0.83037 | 0.80492 | 0.83526 | 0.82203 | 5.6731
E55 | 57.005 | 8.992 | 0.999 | 0.26561 | 0.16757 | 0.074464 | 0.16921 | 0.66343 | 0.63697 | 0.66448 | 0.66107 | 0.65658 | 0.046389
E56 | 5222.568 | 1375.904 | 345.933 | 24.3339 | 25.6402 | 25.7853 | 25.2532 | 0.80889 | 0.88096 | 0.85646 | 0.84276 | 0.84767 | 5.3624
E57 | 5406.863 | 1427.202 | 363.902 | 25.1926 | 26.5962 | 27.1247 | 26.3045 | 0.88979 | 0.9086 | 0.90395 | 0.91296 | 0.90387 | 5.2384
E58 | 51.04 | 6.99 | 0 | 0.23781 | 0.13026 | 0 | 0.12269 | 0.58839 | 0.60181 | 0.59966 | 0.60729 | 0.59933 | 0.036849
E59 | 4460.253 | 1319.268 | 336.03 | 20.782 | 24.5848 | 25.0472 | 23.4713 | 0.9049 | 0.9037 | 0.89062 | 0.916 | 0.90385 | 4.6743
E60 | 54.99 | 10 | 0 | 0.25622 | 0.18635 | 0 | 0.14752 | 0.66063 | 0.66838 | 0.67876 | 0.68949 | 0.6744 | 0.039375
E61 | 444.843 | 106.026 | 25.998 | 2.0727 | 1.9758 | 1.9379 | 1.9955 | 2.237 | 1.8284 | 1.8479 | 2.1137 | 2.0143 | 0.17832
E62 | 49.989 | 6 | 0.999 | 0.23292 | 0.11181 | 0.074464 | 0.13973 | 0.62201 | 0.63512 | 0.65178 | 0.62754 | 0.63421 | 0.039658
E63 | 52.02 | 3.008 | 2.002 | 0.24238 | 0.056055 | 0.14923 | 0.14922 | 0.63859 | 0.63939 | 0.66042 | 0.66845 | 0.65184 | 0.041206
E64 | 352.026 | 76.076 | 20.996 | 1.6402 | 1.4177 | 1.565 | 1.541 | 1.5801 | 1.5326 | 1.5193 | 1.8719 | 1.6323 | 0.16993
F1 | 844.928 | 216 | 58.03 | 3.9368 | 4.0252 | 4.3255 | 4.0958 | 1.4358 | 1.4737 | 1.1641 | 1.2317 | 1.3328 | 0.55316
F2 | 2444.808 | 605.16 | 156.978 | 11.3913 | 11.2773 | 11.7009 | 11.4565 | 0.67037 | 0.66486 | 0.65022 | 0.67504 | 0.66519 | 3.1001
F3 | 131.918 | 14 | 2 | 0.61466 | 0.26089 | 0.14908 | 0.34154 | 0.69345 | 0.70207 | 0.7044 | 0.67736 | 0.6944 | 0.088533
F4 | 137.936 | 14.014 | 2.002 | 0.6427 | 0.26115 | 0.14923 | 0.35103 | 0.76691 | 0.75805 | 0.74184 | 0.73663 | 0.75095 | 0.084139
F5 | 5983.45 | 1501.302 | 378.852 | 27.8792 | 27.9771 | 28.239 | 28.0318 | 0.81963 | 0.85131 | 0.79635 | 0.83077 | 0.82476 | 6.1178
F6 | 5967.626 | 1496.215 | 378.097 | 27.8054 | 27.8823 | 28.1828 | 27.9568 | 0.66565 | 0.67864 | 0.70523 | 0.68787 | 0.6845 | 7.3517
F7 | 6058.8 | 1523.928 | 387.89 | 28.2303 | 28.3987 | 28.9127 | 28.5139 | 0.89085 | 0.89239 | 0.91048 | 0.86584 | 0.89003 | 5.7667
F8 | 5894.468 | 1483.461 | 374.99 | 27.4646 | 27.6446 | 27.9512 | 27.6868 | 0.73158 | 0.75291 | 0.74515 | 0.72449 | 0.73862 | 6.7472
F9 | 136.136 | 13.002 | 1 | 0.63431 | 0.24229 | 0.074538 | 0.31705 | 0.67087 | 0.67213 | 0.67194 | 0.64378 | 0.66479 | 0.085845
F10 | 137.025 | 14.007 | 4 | 0.63845 | 0.26102 | 0.29815 | 0.39921 | 0.61581 | 0.62042 | 0.58324 | 0.6021 | 0.60557 | 0.11866
F11 | 6019.054 | 1540.132 | 384.747 | 28.0451 | 28.7007 | 28.6784 | 28.4747 | 0.69629 | 0.6707 | 0.65608 | 0.69273 | 0.67915 | 7.5469
F12 | 6008.028 | 1541.928 | 384.912 | 27.9937 | 28.7341 | 28.6907 | 28.4729 | 0.73586 | 0.73913 | 0.77035 | 0.76759 | 0.7534 | 6.8027
F13 | 6053.669 | 1563.195 | 387.828 | 28.2064 | 29.1304 | 28.9081 | 28.7483 | 0.75239 | 0.77461 | 0.74475 | 0.71067 | 0.74596 | 6.937
F14 | 5944.796 | 1531.258 | 379.052 | 27.6991 | 28.5353 | 28.2539 | 28.1628 | 0.72482 | 0.71327 | 0.6939 | 0.72061 | 0.71325 | 7.1074
F15 | 5863.767 | 1522.248 | 374.016 | 27.3215 | 28.3674 | 27.8786 | 27.8558 | 0.82892 | 0.86231 | 0.82408 | 0.82836 | 0.83606 | 5.9973
F16 | 5745.327 | 1474.545 | 362.23 | 26.7697 | 27.4784 | 27.0001 | 27.0827 | 0.71806 | 0.75262 | 0.73273 | 0.76633 | 0.74266 | 6.5641
F17 | 5549.94 | 1499.253 | 364.8 | 25.8593 | 27.9389 | 27.1916 | 26.9966 | 0.77716 | 0.79266 | 0.75394 | 0.73355 | 0.76466 | 6.355
F18 | 143.988 | 15.019 | 1.992 | 0.6709 | 0.27988 | 0.14848 | 0.36642 | 0.77667 | 0.77578 | 0.76721 | 0.76296 | 0.77068 | 0.085581
F19 | 145.068 | 16.008 | 1.992 | 0.67593 | 0.29831 | 0.14848 | 0.37424 | 0.91114 | 0.88484 | 0.8805 | 0.91208 | 0.89726 | 0.075077
F20 | 5648.601 | 1467.676 | 358.192 | 26.319 | 27.3504 | 26.6991 | 26.7895 | 0.7277 | 0.7571 | 0.7505 | 0.7256 | 0.74035 | 6.5132
F21 | 5574.576 | 1502.214 | 365.772 | 25.9741 | 27.994 | 27.2641 | 27.0774 | 0.74876 | 0.78743 | 0.78721 | 0.77581 | 0.77496 | 6.2892
F22 | 5600.238 | 1467.669 | 355.992 | 26.0936 | 27.3503 | 26.5351 | 26.6597 | 0.72635 | 0.72717 | 0.71094 | 0.71985 | 0.72111 | 6.6547
F23 | 5656.3 | 1496.292 | 362.929 | 26.3549 | 27.8837 | 27.0522 | 27.0969 | 0.77592 | 0.75612 | 0.74646 | 0.76715 | 0.76149 | 6.4051
F24 | 5599.494 | 1469.585 | 356.96 | 26.0902 | 27.386 | 26.6072 | 26.6945 | 0.80672 | 0.79241 | 0.7961 | 0.80501 | 0.80008 | 6.0056
F25 | 5626.421 | 1493.26 | 364.344 | 26.2156 | 27.8272 | 27.1576 | 27.0668 | 0.82287 | 0.81886 | 0.80833 | 0.828 | 0.81955 | 5.9448
F26 | 5576.816 | 1460.755 | 357.216 | 25.9845 | 27.2215 | 26.6263 | 26.6108 | 0.7378 | 0.73613 | 0.68081 | 0.70251 | 0.71471 | 6.7019
F27 | 5505.763 | 1460.266 | 359.037 | 25.6535 | 27.2123 | 26.7621 | 26.5426 | 0.79359 | 0.77915 | 0.7572 | 0.81415 | 0.7863 | 6.0762
F28 | 5526.53 | 1446.942 | 354.05 | 25.7502 | 26.964 | 26.3903 | 26.3682 | 0.89627 | 0.88291 | 0.86521 | 0.89801 | 0.8857 | 5.3588
F29 | 5620.487 | 1497.2 | 368.98 | 26.188 | 27.9006 | 27.5032 | 27.1973 | 0.83455 | 0.80982 | 0.80472 | 0.78396 | 0.80846 | 6.0553
F30 | 5742.464 | 1492.344 | 367.952 | 26.7563 | 27.8101 | 27.4266 | 27.331 | 0.7306 | 0.71562 | 0.69573 | 0.69215 | 0.70869 | 6.9418
F31 | 5345.592 | 1481.44 | 366.821 | 24.9072 | 27.6069 | 27.3423 | 26.6188 | 0.76711 | 0.73907 | 0.71671 | 0.7565 | 0.74509 | 6.4306
F32 | 5601.225 | 1461.747 | 362.97 | 26.0982 | 27.2399 | 27.0552 | 26.7978 | 0.81027 | 0.74749 | 0.73729 | 0.75366 | 0.76271 | 6.3243
F33 | 105.968 | 58.029 | 18.988 | 0.49375 | 1.0814 | 1.4153 | 0.99682 | 3.1331 | 2.8763 | 2.8438 | 2.8384 | 2.9255 | 0.061333
F34 | 138.88 | 111.111 | 16.074 | 0.64709 | 2.0706 | 1.1981 | 1.3053 | 7.0404 | 7.6423 | 8.2656 | 8.3292 | 7.8369 | 0.02998
F35 | 137.997 | 62.031 | 15.996 | 0.64298 | 1.156 | 1.1923 | 0.99709 | 2.8841 | 2.6827 | 2.5303 | 2.5696 | 2.6702 | 0.067214
F36 | 158.015 | 51.008 | 12.006 | 0.73625 | 0.95054 | 0.89491 | 0.86057 | 0.98188 | 0.97724 | 0.94002 | 0.9629 | 0.96565 | 0.16041
F37 | 160.025 | 48.007 | 12 | 0.74562 | 0.89462 | 0.89446 | 0.8449 | 0.87518 | 0.86708 | 0.84996 | 0.85475 | 0.8618 | 0.17647
F38 | 179.133 | 72.01 | 20.984 | 0.83465 | 1.3419 | 1.5641 | 1.2469 | 2.6258 | 2.4125 | 2.256 | 2.3328 | 2.4107 | 0.093101
F39 | 206.052 | 80 | 22.017 | 0.96007 | 1.4908 | 1.6411 | 1.364 | 2.3687 | 2.2564 | 2.046 | 2.2068 | 2.2225 | 0.11047
F40 | 240.106 | 89.985 | 25.036 | 1.1187 | 1.6769 | 1.8661 | 1.5539 | 2.2713 | 2.1383 | 2.0225 | 2.0737 | 2.1285 | 0.13141
F41 | 196.092 | 81.911 | 20.008 | 0.91367 | 1.5264 | 1.4914 | 1.3105 | 2.2403 | 2.0656 | 1.9246 | 1.9525 | 2.0495 | 0.11509
F42 | 5679.024 | 1485.38 | 364.21 | 26.4607 | 27.6803 | 27.1476 | 27.0962 | 0.73257 | 0.76969 | 0.76653 | 0.80384 | 0.76857 | 6.3459
F43 | 5992.077 | 1573.045 | 395.038 | 27.9194 | 29.314 | 29.4455 | 28.893 | 0.77421 | 0.77085 | 0.77076 | 0.7645 | 0.77009 | 6.7534
F44 | 5802.573 | 1519.855 | 381.996 | 27.0364 | 28.3228 | 28.4734 | 27.9442 | 0.82693 | 0.79035 | 0.77569 | 0.82146 | 0.80389 | 6.257
F45 | 260.15 | 85.07 | 24.975 | 1.2121 | 1.5853 | 1.8616 | 1.553 | 1.6904 | 1.5922 | 1.5867 | 1.6215 | 1.6232 | 0.17222
F46 | 5734.352 | 1493.26 | 377.13 | 26.7185 | 27.8272 | 28.1107 | 27.5521 | 0.77143 | 0.80042 | 0.81446 | 0.8001 | 0.79676 | 6.2245
F47 | 5917.132 | 1546.848 | 393.156 | 27.5702 | 28.8258 | 29.3052 | 28.5671 | 0.71788 | 0.73953 | 0.74996 | 0.7714 | 0.74494 | 6.9026
F48 | 5692.856 | 1485.38 | 374.229 | 26.5252 | 27.6803 | 27.8944 | 27.3667 | 0.88824 | 0.92453 | 0.937 | 0.89252 | 0.91081 | 5.4084
F49 | 5905.221 | 1555.112 | 393.092 | 27.5147 | 28.9798 | 29.3005 | 28.5983 | 0.79642 | 0.81908 | 0.82227 | 0.79256 | 0.80769 | 6.3733
F50 | 5787.628 | 1499.253 | 381.965 | 26.9668 | 27.9389 | 28.4711 | 27.7922 | 0.79935 | 0.78372 | 0.76598 | 0.78077 | 0.78255 | 6.3927
F51 | 5903.3 | 1550.978 | 396.149 | 27.5057 | 28.9028 | 29.5283 | 28.6456 | 0.76804 | 0.77396 | 0.77563 | 0.74734 | 0.76633 | 6.7285
F52 | 5695.82 | 1477.539 | 376.145 | 26.539 | 27.5342 | 28.0373 | 27.3702 | 0.72762 | 0.75471 | 0.7625 | 0.76941 | 0.75373 | 6.5363
F53 | 198.072 | 67.956 | 19.024 | 0.92289 | 1.2664 | 1.418 | 1.2024 | 1.462 | 1.4504 | 1.437 | 1.4744 | 1.456 | 0.14865
F54 | 5699.772 | 1475.565 | 376.145 | 26.5574 | 27.4974 | 28.0373 | 27.364 | 0.75997 | 0.75243 | 0.74071 | 0.72798 | 0.74537 | 6.6082
F55 | 5946.857 | 1564.2 | 396.88 | 27.7087 | 29.1492 | 29.5828 | 28.8136 | 0.75927 | 0.72887 | 0.74977 | 0.73947 | 0.74443 | 6.967
F56 | 5742.99 | 1488.445 | 379.076 | 26.7588 | 27.7375 | 28.2557 | 27.584 | 0.78167 | 0.7775 | 0.76375 | 0.76556 | 0.77216 | 6.4302
F57 | 5920.154 | 1552.73 | 395.038 | 27.5843 | 28.9354 | 29.4455 | 28.6551 | 0.7966 | 0.75267 | 0.75717 | 0.77601 | 0.77081 | 6.6915
F58 | 5848.92 | 1518.115 | 383.756 | 27.2524 | 28.2904 | 28.6046 | 28.0491 | 0.7709 | 0.73488 | 0.73363 | 0.75297 | 0.74825 | 6.7475
F59 | 5946 | 1543.896 | 392.119 | 27.7047 | 28.7708 | 29.2279 | 28.5678 | 0.83117 | 0.84815 | 0.84895 | 0.84451 | 0.84323 | 6.0983
F60 | 5966.637 | 1530.104 | 382.876 | 27.8008 | 28.5138 | 28.539 | 28.2845 | 0.7983 | 0.7952 | 0.78665 | 0.77491 | 0.78882 | 6.4543
F61 | 6290.868 | 1582.161 | 398.132 | 29.3116 | 29.4839 | 29.6761 | 29.4905 | 0.96261 | 0.96766 | 0.93714 | 0.95941 | 0.95678 | 5.5481
F62 | 6557.07 | 1624.896 | 403.914 | 30.5519 | 30.2803 | 30.1071 | 30.3131 | 0.70705 | 0.73295 | 0.7461 | 0.73212 | 0.72969 | 7.4776
F63 | 6535.296 | 1675.926 | 422.064 | 30.4504 | 31.2312 | 31.46 | 31.0472 | 0.73028 | 0.72146 | 0.74276 | 0.72165 | 0.72909 | 7.6651
F64 | 4830.21 | 1274.13 | 326.318 | 22.5058 | 23.7437 | 24.3232 | 23.5242 | 0.78196 | 0.77842 | 0.76674 | 0.75062 | 0.76953 | 5.5025
G1 | 2461.621 | 578.2 | 148.051 | 11.4696 | 10.7749 | 11.0355 | 11.0933 | 0.71807 | 0.73347 | 0.71683 | 0.70951 | 0.71952 | 2.7752
G2 | 2655.465 | 639.265 | 168 | 12.3728 | 11.9128 | 12.5225 | 12.2694 | 0.65985 | 0.66847 | 0.66037 | 0.64869 | 0.65938 | 3.3493
G3 | 28.98 | 0 | 0.999 | 0.13503 | 0 | 0.074464 | 0.069831 | 0.64246 | 0.59536 | 0.57526 | 0.56644 | 0.59561 | 0.021104
G4 | 26.016 | 1.001 | 1 | 0.12122 | 0.018654 | 0.074538 | 0.07147 | 0.72611 | 0.71074 | 0.66049 | 0.68464 | 0.69595 | 0.018485
G5 | 6037.02 | 1486.422 | 381.175 | 28.1288 | 27.6998 | 28.4122 | 28.0802 | 1.1091 | 1.0586 | 1.0588 | 1.0541 | 1.0704 | 4.7222
G6 | 5905.319 | 1499.253 | 387.105 | 27.5151 | 27.9389 | 28.8542 | 28.1027 | 0.87068 | 0.81175 | 0.85021 | 0.85526 | 0.84725 | 5.9705
G7 | 6116.452 | 1507.149 | 382.782 | 28.4989 | 28.086 | 28.532 | 28.3723 | 0.71592 | 0.71506 | 0.66004 | 0.67606 | 0.6922 | 7.378
G8 | 5964.659 | 1516.032 | 384.693 | 27.7916 | 28.2516 | 28.6744 | 28.2392 | 0.77611 | 0.74009 | 0.69589 | 0.72134 | 0.73394 | 6.9257
G9 | 24.012 | 2 | 3.003 | 0.11188 | 0.03727 | 0.22384 | 0.12433 | 0.89777 | 0.80238 | 0.8224 | 0.8246 | 0.83757 | 0.026719
G10 | 18.004 | 2.002 | 1.001 | 0.083888 | 0.037308 | 0.074613 | 0.065269 | 0.73742 | 0.65325 | 0.66375 | 0.66554 | 0.68081 | 0.017257
G11 | 6216.543 | 1574.642 | 397.94 | 28.9652 | 29.3438 | 29.6618 | 29.3236 | 0.75287 | 0.76142 | 0.70032 | 0.74177 | 0.73947 | 7.1379
G12 | 6099.163 | 1569.712 | 391.23 | 28.4183 | 29.2519 | 29.1617 | 28.944 | 0.73426 | 0.69767 | 0.69432 | 0.70958 | 0.70913 | 7.3469
G13 | 6249.491 | 1615.037 | 402.112 | 29.1188 | 30.0965 | 29.9728 | 29.7294 | 0.79247 | 0.82023 | 0.7908 | 0.79029 | 0.79854 | 6.7013
G14 | 6098.924 | 1572.67 | 392.616 | 28.4172 | 29.307 | 29.265 | 28.9964 | 0.789 | 0.80796 | 0.78914 | 0.79146 | 0.79443 | 6.5699
G15 | 6130.811 | 1563.796 | 387.869 | 28.5658 | 29.1416 | 28.9112 | 28.8729 | 0.78273 | 0.76983 | 0.7171 | 0.74377 | 0.75378 | 6.8947
G16 | 6005.895 | 1566.15 | 387.03 | 27.9838 | 29.1855 | 28.8486 | 28.6726 | 0.82031 | 0.76785 | 0.77544 | 0.75539 | 0.78013 | 6.6156
G17 | 6130.08 | 1575.628 | 388 | 28.5624 | 29.3621 | 28.9209 | 28.9485 | 0.72886 | 0.72391 | 0.71196 | 0.71736 | 0.72055 | 7.2316
G18 | 10.992 | 5.005 | 0 | 0.051216 | 0.093269 | 0 | 0.048162 | 0.649 | 0.58913 | 0.57829 | 0.5838 | 0.60073 | 0.014431
G19 | 10.008 | 3 | 1.001 | 0.046631 | 0.055906 | 0.074613 | 0.05905 | 0.65586 | 0.61357 | 0.58308 | 0.59427 | 0.61232 | 0.017358
G20 | 5700.066 | 1499.17 | 365.096 | 26.5588 | 27.9373 | 27.2137 | 27.2366 | 0.79151 | 0.76519 | 0.80133 | 0.82578 | 0.79625 | 6.1571
G21 | 5887.492 | 1540.132 | 372.864 | 27.4321 | 28.7007 | 27.7927 | 27.9751 | 0.90206 | 0.8887 | 0.88248 | 0.86094 | 0.88367 | 5.6984
G22 | 5756.088 | 1516.344 | 368.928 | 26.8198 | 28.2574 | 27.4993 | 27.5255 | 0.74194 | 0.73242 | 0.72267 | 0.69924 | 0.72424 | 6.8411
G23 | 5686.928 | 1491.357 | 360.272 | 26.4976 | 27.7917 | 26.8541 | 27.0478 | 0.89139 | 0.87033 | 0.84133 | 0.88487 | 0.87219 | 5.582
G24 | 5512.395 | 1458.785 | 355.846 | 25.6844 | 27.1847 | 26.5242 | 26.4644 | 0.83882 | 0.80999 | 0.79359 | 0.82831 | 0.81786 | 5.8245
G25 | 5814.38 | 1524.356 | 369.904 | 27.0914 | 28.4067 | 27.5721 | 27.6901 | 0.78134 | 0.76996 | 0.79244 | 0.75634 | 0.77514 | 6.4301
G26 | 5598.996 | 1487.35 | 363.66 | 26.0879 | 27.7171 | 27.1067 | 26.9705 | 0.76562 | 0.77654 | 0.77052 | 0.78381 | 0.77415 | 6.271
G27 | 5829.2 | 1533.798 | 374.957 | 27.1605 | 28.5826 | 27.9487 | 27.8973 | 0.79647 | 0.79627 | 0.77853 | 0.78615 | 0.78939 | 6.3612
G28 | 5529.174 | 1470.63 | 363.951 | 25.7625 | 27.4055 | 27.1283 | 26.7655 | 0.83691 | 0.76589 | 0.73929 | 0.75577 | 0.77536 | 6.2136
G29 | 5832.164 | 1526.027 | 376.915 | 27.1743 | 28.4378 | 28.0947 | 27.9022 | 0.85696 | 0.79864 | 0.8024 | 0.7931 | 0.81318 | 6.1762
G30 | 5514.369 | 1465.695 | 364.188 | 25.6936 | 27.3135 | 27.146 | 26.7177 | 0.72184 | 0.68798 | 0.7157 | 0.71145 | 0.70936 | 6.7796
G31 | 3 | 0.999 | 0 | 0.013978 | 0.018617 | 0 | 0.010865 | 0.31847 | 0.32698 | 0.31194 | 0.31151 | 0.31729 | 0.0061638
G32 | 5546.632 | 1467.65 | 364.875 | 25.8439 | 27.3499 | 27.1972 | 26.797 | 0.8388 | 0.78281 | 0.75271 | 0.77464 | 0.78788 | 6.1221
G33 | 5642.01 | 1479.544 | 362.88 | 26.2883 | 27.5716 | 27.0485 | 26.9695 | 0.81081 | 0.78534 | 0.81169 | 0.83426 | 0.81071 | 5.988
G34 | 5500.894 | 1461.742 | 365.024 | 25.6308 | 27.2398 | 27.2083 | 26.693 | 0.76147 | 0.79571 | 0.78439 | 0.83356 | 0.79421 | 6.0497
G35 | 5609.608 | 1471.551 | 362.068 | 26.1373 | 27.4226 | 26.988 | 26.8493 | 0.92622 | 0.94916 | 0.9415 | 0.97357 | 0.94777 | 5.0992
G36 | 111.93 | 8.987 | 1.001 | 0.52152 | 0.16747 | 0.074613 | 0.25454 | 0.87202 | 0.84068 | 0.80888 | 0.83167 | 0.83862 | 0.054634
G37 | 119.988 | 10.999 | 0 | 0.55907 | 0.20497 | 0 | 0.25468 | 0.90533 | 0.83521 | 0.82216 | 0.84625 | 0.85283 | 0.053753
G38 | 125.028 | 10.005 | 1.001 | 0.58255 | 0.18645 | 0.074613 | 0.2812 | 0.97411 | 0.92656 | 0.87205 | 0.86086 | 0.90953 | 0.055652
G39 | 11411.082 | 2991.725 | 740.722 | 53.1686 | 55.7514 | 55.2123 | 54.7107 | 2.0637 | 2.0413 | 2.0901 | 2.1375 | 2.0834 | 4.7268
G40 | 5580.025 | 1474.443 | 370.713 | 25.9995 | 27.4765 | 27.6324 | 27.0361 | 0.95199 | 0.96916 | 0.90468 | 0.94568 | 0.94318 | 5.1597
G41 | 133.056 | 11.001 | 2.002 | 0.61996 | 0.20501 | 0.14923 | 0.32473 | 0.96473 | 0.90338 | 0.8875 | 0.89854 | 0.91403 | 0.063949
G42 | 143.929 | 14.01 | 3.006 | 0.67062 | 0.26108 | 0.22406 | 0.38525 | 1.9148 | 1.8423 | 1.757 | 1.8204 | 1.8345 | 0.037801
G43 | 134.976 | 11 | 1.001 | 0.6289 | 0.20499 | 0.074613 | 0.30283 | 0.94891 | 0.91247 | 0.93314 | 0.8958 | 0.9228 | 0.05907
G44 | 87.971 | 8.001 | 2.006 | 0.40989 | 0.1491 | 0.14952 | 0.23617 | 1.5468 | 1.4864 | 1.4036 | 1.4293 | 1.4676 | 0.028967
G45 | 84.051 | 8 | 3 | 0.39163 | 0.14908 | 0.22362 | 0.25477 | 1.4068 | 1.4505 | 1.3936 | 1.4009 | 1.4132 | 0.032452
G46 | 95.076 | 7.002 | 2.006 | 0.443 | 0.13048 | 0.14952 | 0.241 | 1.5068 | 1.4759 | 1.4927 | 1.472 | 1.4869 | 0.029174
G47 | 129.129 | 0.994 | 0 | 0.60166 | 0.018523 | 0 | 0.20673 | 1.5608 | 1.6086 | 1.529 | 1.6267 | 1.5818 | 0.023525
G48 | 148.96 | 11.991 | 1.001 | 0.69406 | 0.22345 | 0.074613 | 0.33071 | 1.079 | 1.0548 | 1.0182 | 1.0405 | 1.0483 | 0.056782
G49 | 175.89 | 11 | 0.992 | 0.81954 | 0.20499 | 0.073942 | 0.36616 | 1.8186 | 1.6941 | 1.6362 | 1.6728 | 1.7068 | 0.038615
G50 | 178.012 | 10.01 | 3.003 | 0.82943 | 0.18654 | 0.22384 | 0.41327 | 1.7009 | 1.6098 | 1.6444 | 1.5801 | 1.6344 | 0.045513
G51 | 183.134 | 12 | 1.005 | 0.85329 | 0.22362 | 0.074911 | 0.38394 | 1.7458 | 1.7106 | 1.57 | 1.6289 | 1.6652 | 0.041502
G52 | 212.054 | 15.996 | 0.99 | 0.98804 | 0.29809 | 0.073793 | 0.45331 | 1.8944 | 1.887 | 1.7266 | 1.8322 | 1.8363 | 0.044435
G53 | 235.116 | 20.026 | 3.99 | 1.0955 | 0.37319 | 0.29741 | 0.5887 | 1.9304 | 1.8295 | 1.8066 | 1.7759 | 1.8365 | 0.057699
G54 | 267.036 | 28 | 3.002 | 1.2442 | 0.52179 | 0.22376 | 0.66326 | 2.0406 | 1.9494 | 1.8723 | 1.9072 | 1.9434 | 0.061432
G55 | 269.99 | 25 | 0 | 1.258 | 0.46588 | 0 | 0.57462 | 1.9994 | 1.9261 | 1.7815 | 1.8563 | 1.8926 | 0.054652
G56 | 289.014 | 31.988 | 4.002 | 1.3466 | 0.5961 | 0.2983 | 0.74701 | 2.0677 | 1.9632 | 1.9257 | 1.8773 | 1.9597 | 0.068612
G57 | 317.245 | 34.986 | 3 | 1.4782 | 0.65197 | 0.22362 | 0.78458 | 2.1785 | 2.0671 | 1.9654 | 1.9991 | 2.0541 | 0.068752
G58 | 336.96 | 37.98 | 6.984 | 1.57 | 0.70776 | 0.52058 | 0.93279 | 2.272 | 2.1347 | 2.0532 | 2.0727 | 2.1349 | 0.078648
G59 | 360.045 | 46.02 | 5.017 | 1.6776 | 0.85759 | 0.37396 | 0.96971 | 2.3769 | 2.2097 | 2.1554 | 2.1851 | 2.2334 | 0.078154
G60 | 165.977 | 14.007 | 2.002 | 0.77335 | 0.26102 | 0.14923 | 0.39453 | 1.1391 | 1.0554 | 1.0208 | 1.064 | 1.0707 | 0.066327
G61 | 364.941 | 50.022 | 7.006 | 1.7004 | 0.93217 | 0.52222 | 1.0516 | 2.4199 | 2.2957 | 2.1236 | 2.2272 | 2.2691 | 0.083418
G62 | 407.886 | 50.991 | 3.99 | 1.9005 | 0.95023 | 0.29741 | 1.0494 | 2.5601 | 2.4343 | 2.2432 | 2.4173 | 2.4164 | 0.07817
G63 | 443.196 | 56.056 | 9 | 2.065 | 1.0446 | 0.67085 | 1.2602 | 2.7556 | 2.5583 | 2.481 | 2.5288 | 2.583 | 0.087815
G64 | 503.965 | 72.03 | 18.001 | 2.3482 | 1.3423 | 1.3418 | 1.6774 | 3.3166 | 3.0669 | 2.9727 | 3.0557 | 3.1057 | 0.097221
H1 | 2517.14 | 589.645 | 152.154 | 11.7283 | 10.9881 | 11.3413 | 11.3526 | 0.66774 | 0.66822 | 0.64242 | 0.66978 | 0.66214 | 3.0862
H2 | 2667.772 | 633.696 | 166.041 | 12.4302 | 11.809 | 12.3764 | 12.2052 | 0.62315 | 0.61169 | 0.6432 | 0.62655 | 0.62625 | 3.5081
H3 | 12.006 | 9 | 1 | 0.055941 | 0.16772 | 0.074538 | 0.099399 | 0.71925 | 0.68222 | 0.6182 | 0.65148 | 0.66883 | 0.026751
H4 | 15.988 | 8.993 | 4 | 0.074494 | 0.16759 | 0.29815 | 0.18008 | 0.67588 | 0.66837 | 0.65607 | 0.65235 | 0.66323 | 0.048873
H5 | 4901.486 | 1199.205 | 312 | 22.8379 | 22.3474 | 23.256 | 22.8138 | 0.69184 | 0.69939 | 0.70467 | 0.70258 | 0.69964 | 5.8694
H6 | 4770.936 | 1179.816 | 307.173 | 22.2296 | 21.9861 | 22.8962 | 22.3706 | 0.74205 | 0.75991 | 0.77808 | 0.74759 | 0.75703 | 5.3191
H7 | 4843.017 | 1196.018 | 307.188 | 22.5655 | 22.288 | 22.8973 | 22.5836 | 0.70658 | 0.73115 | 0.73896 | 0.71679 | 0.72348 | 5.6188
H8 | 4782.69 | 1196.018 | 308.1 | 22.2844 | 22.288 | 22.9653 | 22.5126 | 0.73599 | 0.72304 | 0.71684 | 0.70244 | 0.71968 | 5.6306
H9 | 6 | 4.998 | 1.998 | 0.027956 | 0.093139 | 0.14893 | 0.090008 | 0.64422 | 0.60428 | 0.60213 | 0.60896 | 0.61513 | 0.026338
H10 | 7.008 | 6 | 0.999 | 0.032653 | 0.11181 | 0.074464 | 0.072976 | 0.5851 | 0.54249 | 0.522 | 0.53594 | 0.54689 | 0.024019
H11 | 5935.161 | 1499.383 | 376.915 | 27.6542 | 27.9413 | 28.0947 | 27.8967 | 0.69241 | 0.70413 | 0.67109 | 0.67127 | 0.68487 | 7.3319
H12 | 5783.58 | 1471.551 | 370.062 | 26.9479 | 27.4226 | 27.5838 | 27.3181 | 0.68984 | 0.70559 | 0.70227 | 0.70666 | 0.70112 | 7.0134
H13 | 5488.95 | 1481.487 | 373.214 | 25.5751 | 27.6078 | 27.8188 | 27.0006 | 0.72308 | 0.72241 | 0.72571 | 0.73579 | 0.72677 | 6.6873
H14 | 5878.616 | 1507.149 | 373.808 | 27.3907 | 28.086 | 27.8631 | 27.7799 | 0.69938 | 0.72265 | 0.69948 | 0.71224 | 0.7085 | 7.0577
H15 | 5772.575 | 1479.513 | 365.848 | 26.8966 | 27.571 | 27.2697 | 27.2458 | 0.86598 | 0.84161 | 0.85515 | 0.82675 | 0.8475 | 5.7867
H16 | 5737.89 | 1479.513 | 367.952 | 26.735 | 27.571 | 27.4266 | 27.2442 | 0.71699 | 0.75429 | 0.7593 | 0.73862 | 0.74248 | 6.6048
H17 | 5773.566 | 1488.396 | 366.894 | 26.9012 | 27.7365 | 27.3477 | 27.3285 | 0.71412 | 0.67849 | 0.68751 | 0.68212 | 0.6907 | 7.122
H18 | 0 | 5.005 | 0 | 0 | 0.093269 | 0 | 0.03109 | 0.61352 | 0.56445 | 0.56977 | 0.57426 | 0.58082 | 0.0096348
H19 | 2.002 | 2.002 | 0 | 0.0093281 | 0.037308 | 0 | 0.015545 | 0.59687 | 0.57972 | 0.57387 | 0.58675 | 0.58437 | 0.0047883
H20 | 5746.809 | 1492.401 | 365.024 | 26.7766 | 27.8112 | 27.2083 | 27.2654 | 0.7278 | 0.72271 | 0.74894 | 0.76398 | 0.74104 | 6.6228
H21 | 5719.061 | 1492.401 | 366.146 | 26.6473 | 27.8112 | 27.292 | 27.2501 | 0.77438 | 0.76051 | 0.76272 | 0.79685 | 0.77375 | 6.3393
H22 | 5526.532 | 1448.935 | 354.9 | 25.7502 | 27.0012 | 26.4537 | 26.4017 | 0.75047 | 0.78604 | 0.75375 | 0.7743 | 0.76628 | 6.2018
H23 | 5521.23 | 1476.552 | 358.16 | 25.7255 | 27.5158 | 26.6967 | 26.646 | 0.73459 | 0.7504 | 0.75286 | 0.71489 | 0.73834 | 6.496
H24 | 5517.631 | 1449.903 | 351.864 | 25.7087 | 27.0192 | 26.2274 | 26.3185 | 0.70007 | 0.72494 | 0.67642 | 0.69757 | 0.69996 | 6.768
H25 | 5596.47 | 1465.698 | 356.118 | 26.0761 | 27.3136 | 26.5445 | 26.6447 | 0.78179 | 0.78622 | 0.81865 | 0.76278 | 0.78761 | 6.0893
H26 | 5684.772 | 1497.279 | 363.902 | 26.4875 | 27.9021 | 27.1247 | 27.1714 | 0.78095 | 0.75571 | 0.76922 | 0.7499 | 0.76404 | 6.4013
H27 | 5405.4 | 1433.124 | 348.23 | 25.1858 | 26.7065 | 25.9565 | 25.9496 | 0.83825 | 0.83265 | 0.83234 | 0.83693 | 0.83505 | 5.5936
H28 | 5469.568 | 1440.033 | 354.172 | 25.4848 | 26.8353 | 26.3994 | 26.2398 | 0.78461 | 0.80593 | 0.76645 | 0.78685 | 0.78608 | 6.0085
H29 | 5550.93 | 1465.68 | 360.096 | 25.8639 | 27.3132 | 26.841 | 26.6727 | 0.83475 | 0.82525 | 0.81033 | 0.79886 | 0.81742 | 5.8735
H30 | 5533.455 | 1454.838 | 359.168 | 25.7825 | 27.1112 | 26.7718 | 26.5552 | 0.68428 | 0.6954 | 0.71608 | 0.73146 | 0.70704 | 6.7605
H31 | 5453.473 | 1448.885 | 357.335 | 25.4098 | 27.0003 | 26.6352 | 26.3484 | 0.72052 | 0.7112 | 0.71386 | 0.69274 | 0.70966 | 6.6831
H32 | 5586.861 | 1470.643 | 362.929 | 26.0313 | 27.4057 | 27.0522 | 26.8297 | 0.77202 | 0.79059 | 0.78379 | 0.80982 | 0.78917 | 6.1195
H33 | 4754.844 | 1479 | 357.822 | 22.1546 | 27.5615 | 26.6715 | 25.4625 | 1.1926 | 1.2033 | 1.2972 | 1.2783 | 1.2437 | 3.6852
H34 | 5863.781 | 1542.132 | 384.912 | 27.3216 | 28.7379 | 28.6907 | 28.2501 | 1.0085 | 0.97625 | 1.0843 | 1.0371 | 1.0273 | 4.9499
H35 | 5565.456 | 1513.257 | 375.18 | 25.9316 | 28.1998 | 27.9653 | 27.3656 | 1.1106 | 1.0927 | 1.1304 | 1.0801 | 1.1036 | 4.4634
H36 | 31.988 | 1.008 | 5.005 | 0.14904 | 0.018784 | 0.37306 | 0.1803 | 0.78802 | 0.76783 | 0.82473 | 0.81401 | 0.79895 | 0.04062
H37 | 36.006 | 4.012 | 1.008 | 0.16777 | 0.074764 | 0.075135 | 0.10589 | 0.90662 | 0.93853 | 0.97305 | 0.93972 | 0.93977 | 0.020281
H38 | 5846.984 | 1536.6 | 385.726 | 27.2433 | 28.6348 | 28.7514 | 28.2099 | 0.88223 | 0.85467 | 0.88728 | 0.89831 | 0.88077 | 5.7651
H39 | 5693.49 | 1496.357 | 371.14 | 26.5281 | 27.8849 | 27.6642 | 27.3591 | 0.88734 | 0.84459 | 0.87715 | 0.92513 | 0.88402 | 5.5707
H40 | 5607.63 | 1464.67 | 370.062 | 26.1281 | 27.2944 | 27.5838 | 27.0021 | 0.90061 | 0.92226 | 0.91181 | 0.91231 | 0.91178 | 5.3306
H41 | 71.012 | 6.996 | 1.003 | 0.33087 | 0.13037 | 0.074762 | 0.17867 | 0.71288 | 0.71074 | 0.68246 | 0.69366 | 0.70005 | 0.04594
H42 | 5612.575 | 1466.665 | 370.062 | 26.1511 | 27.3316 | 27.5838 | 27.0222 | 0.75964 | 0.7867 | 0.81255 | 0.84538 | 0.8017 | 6.0671
H43 | 5606.784 | 1489.32 | 366.102 | 26.1241 | 27.7538 | 27.2887 | 27.0555 | 0.76201 | 0.77557 | 0.78863 | 0.80048 | 0.7818 | 6.2292
H44 | 5671.71 | 1484.518 | 373.042 | 26.4267 | 27.6643 | 27.806 | 27.299 | 0.88011 | 0.83034 | 0.8731 | 0.88351 | 0.86703 | 5.6674
H45 | 5666.76 | 1499.253 | 369.72 | 26.4036 | 27.9389 | 27.5584 | 27.3003 | 0.90683 | 0.86827 | 0.9213 | 0.939 | 0.90922 | 5.4047
H46 | 5745.101 | 1501.227 | 380.034 | 26.7686 | 27.9757 | 28.3271 | 27.6905 | 0.77272 | 0.73281 | 0.74727 | 0.77998 | 0.75843 | 6.5718
H47 | 5874.624 | 1518.87 | 373.906 | 27.3721 | 28.3044 | 27.8704 | 27.849 | 0.81655 | 0.8187 | 0.77362 | 0.85283 | 0.81591 | 6.1439
H48 | 5544 | 1434.197 | 364.125 | 25.8316 | 26.7265 | 27.1413 | 26.5665 | 0.87631 | 0.86805 | 0.87735 | 0.88823 | 0.87752 | 5.4494
H49 | 5527.424 | 1456.812 | 365.148 | 25.7544 | 27.148 | 27.2176 | 26.7066 | 0.91277 | 0.91695 | 0.88546 | 0.8701 | 0.89653 | 5.362
H50 | 5688.54 | 1468.602 | 370.928 | 26.5051 | 27.3677 | 27.6484 | 27.1737 | 0.953 | 0.91389 | 0.89744 | 0.93452 | 0.92495 | 5.2881
H51 | 106.92 | 18.981 | 1.008 | 0.49818 | 0.35371 | 0.075135 | 0.30901 | 0.74865 | 0.70079 | 0.67748 | 0.71748 | 0.71157 | 0.078168
H52 | 5647.95 | 1465.695 | 370.922 | 26.316 | 27.3135 | 27.6479 | 27.0925 | 0.85549 | 0.84191 | 0.85719 | 0.89411 | 0.86239 | 5.6548
H53 | 111.006 | 18 | 2.002 | 0.51722 | 0.33543 | 0.14923 | 0.33396 | 0.80405 | 0.8033 | 0.78716 | 0.79061 | 0.79632 | 0.075489
H54 | 5741.01 | 1493.39 | 382.007 | 26.7496 | 27.8296 | 28.4742 | 27.6845 | 0.83509 | 0.85362 | 0.8931 | 0.9058 | 0.87238 | 5.7122
H55 | 5939.01 | 1531.961 | 386 | 27.6721 | 28.5484 | 28.7718 | 28.3308 | 0.83901 | 0.88505 | 0.84406 | 0.89872 | 0.86709 | 5.8812
H56 | 5807.34 | 1505.175 | 378.805 | 27.0586 | 28.0492 | 28.2355 | 27.7811 | 0.98483 | 0.94275 | 0.97525 | 1.0199 | 0.98106 | 5.0972
H57 | 5822.048 | 1510.203 | 380.952 | 27.1271 | 28.1429 | 28.3956 | 27.8885 | 0.85714 | 0.84005 | 0.80583 | 0.83762 | 0.83536 | 6.0093
H58 | 112.896 | 10.005 | 4.011 | 0.52603 | 0.18645 | 0.29897 | 0.33715 | 0.87907 | 0.85763 | 0.8095 | 0.80954 | 0.83948 | 0.07229
H59 | 5928.162 | 1527.876 | 385.09 | 27.6216 | 28.4723 | 28.704 | 28.2659 | 0.88383 | 0.84729 | 0.85931 | 0.88986 | 0.87025 | 5.8465
H60 | 112.118 | 10.983 | 0 | 0.5224 | 0.20467 | 0 | 0.24236 | 0.82713 | 0.81697 | 0.76318 | 0.78366 | 0.79814 | 0.054657
H61 | 112.1 | 8.987 | 0.996 | 0.52232 | 0.16747 | 0.07424 | 0.25468 | 0.833 | 0.82648 | 0.79554 | 0.82404 | 0.81989 | 0.055912
H62 | 6623.1 | 1590.144 | 400.052 | 30.8596 | 29.6326 | 29.8193 | 30.1038 | 0.8675 | 0.80894 | 0.86255 | 0.8406 | 0.84521 | 6.411
H63 | 6348.346 | 1746.574 | 432.095 | 29.5794 | 32.5477 | 32.2077 | 31.4449 | 0.80152 | 0.80779 | 0.78386 | 0.80454 | 0.79948 | 7.0797
H64 | 6237.99 | 1638.732 | 429.781 | 29.0652 | 30.5381 | 32.0352 | 30.5462 | 0.8867 | 0.9015 | 0.88731 | 0.94054 | 0.90428 | 6.0803

## Pixel Map
Filter | Channel | Center Wavelength | D* | Row | Row | Row | Row | Row | Row | Row
Number | Number | (microns) | Row A | B | C | D | E | F | G | H
F0 | 0 | 0 | 3.7774 | 2.9952 | 1.9071 | 2.6123 | 2.5174 | 0.55316 | 2.7752 | 3.0862
| 1 | 0.84375 | 3.6288 | 3.5594 | 2.0372 | 2.5998 | 1.5676 | 3.1001 | 3.3493 | 3.5081
| 2 | 1.6875 | 0.031809 | 0.038079 | 0.036636 | 0.03542 | 0.027365 | 0.088533 | 0.021104 | 0.026751
| 3 | 2.5312 | 0.020633 | 0.028338 | 0.034511 | 0.035587 | 0.032119 | 0.084139 | 0.018485 | 0.048873
F1 | 4 | 3.375 | 7.7153 | 6.0793 | 5.9176 | 5.9808 | 6.0082 | 6.1178 | 4.7222 | 5.8694
| 5 | 4.2188 | 7.9424 | 6.5109 | 6.5876 | 6.5795 | 6.8311 | 7.3517 | 5.9705 | 5.3191
| 6 | 5.0625 | 7.523 | 5.2889 | 6.7421 | 6.006 | 7.1461 | 5.7667 | 7.378 | 5.6188
| 7 | 5.9062 | 7.0265 | 5.9271 | 6.143 | 5.4177 | 6.3472 | 6.7472 | 6.9257 | 5.6306
| 8 | 6.75 | 0.028314 | 0.027282 | 0.040826 | 0.047158 | 0.040509 | 0.085845 | 0.026719 | 0.026338
| 9 | 7.5938 | 0.018176 | 0.033977 | 0.034276 | 0.034843 | 0.056768 | 0.11866 | 0.017257 | 0.024019
F2 | 10 | 8.4375 | 8.4496 | 6.105 | 6.1066 | 5.5881 | 6.0707 | 7.5469 | 7.1379 | 7.3319
| 11 | 9.2812 | 7.9842 | 6.2103 | 6.4302 | 4.6982 | 6.5632 | 6.8027 | 7.3469 | 7.0134
| 12 | 10.125 | 7.9683 | 6.4312 | 5.9368 | 6.1134 | 6.4569 | 6.937 | 6.7013 | 6.6873
| 13 | 10.9688 | 8.3539 | 6.4241 | 6.3058 | 5.6285 | 6.5463 | 7.1074 | 6.5699 | 7.0577
| 14 | 11.8125 | 8.6904 | 7.0037 | 6.3424 | 6.3229 | 6.1021 | 5.9973 | 6.8947 | 5.7867
| 15 | 12.6562 | 7.778 | 6.5272 | 6.3078 | 6.4566 | 5.912 | 6.5641 | 6.6156 | 6.6048
| 16 | 13.5 | 7.447 | 5.9106 | 6.8116 | 5.9052 | 6.4762 | 6.355 | 7.2316 | 7.122
| 17 | 14.3438 | 0.037005 | 0.02641 | 0.050867 | 0.04306 | 0.048168 | 0.085581 | 0.014431 | 0.0096348
| 18 | 15.1875 | 0.027224 | 0.01939 | 0.041902 | 0.023908 | 0.061656 | 0.075077 | 0.017358 | 0.0047883
F3 | 19 | 16.0312 | 7.6819 | 5.8811 | 6.1032 | 5.3016 | 6.327 | 6.5132 | 6.1571 | 6.6228
| 20 | 16.875 | 8.0596 | 5.9124 | 6.0569 | 6.4791 | 7.0924 | 6.2892 | 5.6984 | 6.3393
| 21 | 17.7188 | 7.1936 | 6.7354 | 6.1504 | 6.4381 | 7.0169 | 6.6547 | 6.8411 | 6.2018
| 22 | 18.5625 | 8.141 | 6.3975 | 5.7722 | 5.8596 | 6.0036 | 6.4051 | 5.582 | 6.496
| 23 | 19.4062 | 8.3605 | 6.3917 | 4.9497 | 5.8373 | 7.0304 | 6.0056 | 5.8245 | 6.768
| 24 | 20.25 | 7.5595 | 5.9712 | 5.4756 | 5.9063 | 6.2395 | 5.9448 | 6.4301 | 6.0893
| 25 | 21.0938 | 8.1693 | 6.7012 | 5.367 | 6.6069 | 7.4741 | 6.7019 | 6.271 | 6.4013
| 26 | 21.9375 | 8.898 | 5.8009 | 5.2427 | 6.3176 | 7.0139 | 6.0762 | 6.3612 | 5.5936
| 27 | 22.7812 | 5.3364 | 6.9043 | 5.419 | 6.4207 | 7.2592 | 5.3588 | 6.2136 | 6.0085
| 28 | 23.625 | 7.4133 | 6.2447 | 4.4875 | 5.4886 | 7.2036 | 6.0553 | 6.1762 | 5.8735
| 29 | 24.4688 | 7.2698 | 6.5191 | 4.7211 | 4.9333 | 6.4879 | 6.9418 | 6.7796 | 6.7605
| 30 | 25.3125 | 8.2858 | 6.757 | 2.4002 | 6.3597 | 7.0227 | 6.4306 | 0.0061638 | 6.6831
| 31 | 26.1562 | 7.5481 | 6.3779 | 4.2376 | 6.14 | 6.2769 | 6.3243 | 6.1221 | 6.1195
| 32 | 27 | 7.0982 | 4.837 | 0.015024 | 6.2526 | 6.4993 | 0.061333 | 5.988 | 3.6852
| 33 | 27.8438 | 7.8589 | 4.8305 | 0.012514 | 5.3861 | 6.1094 | 0.02998 | 6.0497 | 4.9499
| 34 | 28.6875 | 7.6146 | 4.9841 | 0.009493 | 6.0616 | 5.9929 | 0.067214 | 5.0992 | 4.4634
| 35 | 29.5312 | 0.029029 | 0.024465 | 0.0022897 | 0.015824 | 0.019327 | 0.16041 | 0.054634 | 0.04062
| 36 | 30.375 | 0.026837 | 0.011844 | 0.00923 | 0.01967 | 0.023598 | 0.17647 | 0.053753 | 0.020281
F4 | 37 | 31.2188 | 7.7174 | 5.2018 | 0.00045084 | 6.7888 | 5.8278 | 0.093101 | 0.055652 | 5.7651
| 38 | 32.0625 | 0.020741 | 4.3002 | 0.0037281 | 5.9777 | 5.6511 | 0.11047 | 4.7268 | 5.5707
| 39 | 32.9062 | 0.036171 | 4.9708 | 0.00043378 | 5.8646 | 0.035274 | 0.13141 | 5.1597 | 5.3306
| 40 | 33.75 | 7.1526 | 4.5687 | 0.0037269 | 5.7855 | 5.8185 | 0.11509 | 0.063949 | 0.04594
| 41 | 34.5938 | 7.0145 | 4.5399 | 0.0086547 | 6.2352 | 5.6957 | 6.3459 | 0.037801 | 6.0671
| 42 | 35.4375 | 6.6236 | 4.9988 | 0.009229 | 6.4673 | 5.6803 | 6.7534 | 0.05907 | 6.2292
| 43 | 36.2812 | 6.6153 | 5.342 | 0.010416 | 6.6535 | 0.049128 | 6.257 | 0.028967 | 5.6674
| 44 | 37.125 | 6.4955 | 4.5256 | 0.022732 | 6.3498 | 6.3009 | 0.17222 | 0.032452 | 5.4047
| 45 | 37.9688 | 6.2027 | 5.0353 | 0.018592 | 6.1564 | 0.29776 | 6.2245 | 0.029174 | 6.5718
| 46 | 38.8125 | 6.0485 | 4.9527 | 0.0012148 | 6.8431 | 0.053541 | 6.9026 | 0.023525 | 6.1439
| 47 | 39.6562 | 5.6377 | 5.3665 | 0.00083955 | 5.458 | 0.050422 | 5.4084 | 0.056782 | 5.4494
| 48 | 40.5 | 0.45491 | 5.0324 | 0.0041421 | 5.8543 | 6.5966 | 6.3733 | 0.038615 | 5.362
| 49 | 41.3438 | 0.16353 | 5.4181 | 0.010052 | 6.1951 | 0.03112 | 6.3927 | 0.045513 | 5.2881
| 50 | 42.1875 | 5.189 | 5.2094 | 0.011155 | 6.7355 | 0.17404 | 6.7285 | 0.041502 | 0.078168
| 51 | 43.0312 | 5.9897 | 5.0281 | 0.0086505 | 6.4796 | 0.77676 | 6.5363 | 0.044435 | 5.6548
| 52 | 43.875 | 5.5089 | 5.0157 | 0.011006 | 0.080421 | 0.045994 | 0.14865 | 0.057699 | 0.075489
| 53 | 44.7188 | 4.8645 | 5.4132 | 0.016727 | 6.097 | 5.6731 | 6.6082 | 0.061432 | 5.7122
| 54 | 45.5625 | 5.5178 | 4.7556 | 0.010896 | 6.4482 | 0.046389 | 6.967 | 0.054652 | 5.8812
| 55 | 46.4062 | 5.5367 | 5.4922 | 0.013031 | 5.8576 | 5.3624 | 6.4302 | 0.068612 | 5.0972
| 56 | 47.25 | 5.3619 | 5.278 | 0.010404 | 5.1135 | 5.2384 | 6.6915 | 0.068752 | 6.0093
| 57 | 48.0938 | 5.4478 | 5.2912 | 0.013798 | 5.2169 | 0.036849 | 6.7475 | 0.078648 | 0.07229
| 58 | 48.9375 | 5.6419 | 5.8616 | 0.0091623 | 6.6782 | 4.6743 | 6.0983 | 0.078154 | 5.8465
| 59 | 49.7812 | 6.0598 | 3.4208 | 0.00819 | 5.7089 | 0.039375 | 6.4543 | 0.066327 | 0.054657
| 60 | 50.625 | 6.0469 | 5.9822 | 0.011701 | 5.8574 | 0.17832 | 5.5481 | 0.083418 | 0.055912
| 61 | 51.4688 | 6.0964 | 6.2622 | 0.019571 | 6.2651 | 0.039658 | 7.4776 | 0.07817 | 6.411
| 62 | 52.3125 | 6.3525 | 6.6919 | 0.015102 | 7.134 | 0.041206 | 7.6651 | 0.087815 | 7.0797
| 63 | 53.1562 | 0.028023 | 6.2483 | 0.023141 | 4.9254 | 0.16993 | 5.5025 | 0.097221 | 6.0803